## Slide 1: The Electron-Beam Ion Trap (EBIT) of the ReA post-accelerator
A. Lapierre, S. Schwarz, T. M Baumann, K. Kittimanapun, W. Wittmer, C. Sumithrarachchi, A. J. Rodriguez, S. J. Williams, A. Villari, G. Bollen and D. Leitner
National Superconducting Cyclotron Laboratory, Michigan State University, East Lansing, USA

## Slide 2: Outline
  PART I
  Motivation for a post-accelerator: ReA
  A few words on the future Facility for Rare Isotope Beams (FRIB)
  Reacceleration concept
  ReA EBIT charge breeder (key parameters, main charactersitics)
  PART II (some real physics…)
  “Charge breeding” physics
  Physics of (continuous) beam injection
  PART III
  Commissioning results:
    Efficiency, e-beam current density, time structure, contaminations….
  Summary & Outlook

## Slide 3: ReA: Reaccelerator of rare isotopes
ReA, post-accelerator
Beam “stopping”
He gas cells
Acceleration (>80 MeV/u):
Electron Cyclotron Ion Sources + coupled cyclotrons
NSCL: National Superconducting Cyclotron Laboratory
  User facility producing rare-isotope beams by projectile fragmentation.
ReA@NSCL
  Post-accelerator for reacceleration of rare-isotope beams to several MeV/u.
  Why do we need to reach this energy regime ?
  Key reactions in nuclear astrophysics.
  Nuclear structure studies
  e.g., nuclear physics…
EBIT charge breeder
“Low” Energy Beam Area
High Energy Beam Area
Production (>80 MeV/u):
Projectile fragmentation & separation
Target, e.g., beryllium
NSCL rates ≤ 109 pps
How does this entire system work?
In short…
Accel.  Production  Stopping  Reaccel.
High-energy stable-isotope beams from 2 coupled cyclotrons transported to a target where rare-isotope beams are produced at high energies
To produce beams of a few MeV/u, rare isotopes are first thermalized in a gas cell…
(Thermalization)

## Slide 4: Facility for Rare Isotope Beams (FRIB)
In 2008, NSCL & MSU selected to establish the US facility (FRIB) for science with rare-isotope beams  Early completion ~2020
The two NSCL coupled cyclotrons will be later replaced by a 400-kW superconducting heavy-ion “driver” linac.
Coupled cyclotrons

## Slide 5: Facility for Rare Isotope Beams (FRIB)
EBIT charge breeder
Different perspective…
FRIB will increase the current production rates of rare isotopes by up to 2 orders of magnitude < 1012 pps

## Slide 6: The ReAcceleration concept
  After production, rare isotopes are “stopped” in a He gas cell
  Injection into an electron beam ion source/trap for charge breeding (1+  Q+)
  Extraction to Q/A separator and injection into a RFQ, and SRF cavities for reacceleration
Final configuration, ReA (x4)
First configuration, ReA3
48Ca 0.3 - 6 MeV/u
238U 0.3 - 3 MeV/u
A few MeV/u,
Efinal (F, A, Q)
Quasi-continuous injection
Pulsed extrac. rates: 5 – 50 Hz
He gas-cell
Room-temperature RFQ
b=4.1%
EBIT charge breeder
1+  Q+
Trapped ions
< 500 eV
≤ 60 keV
2≤ A/Q ≤ 5
Magnetic
sector
Achromatic
Q/A separator
Electrostatic
sector
Continuous stable heavy ion beam >80 MeV/u
Superconducting RF linac
Thin foil
target
<1 eV
b=8.5%
600 keV/u
80 MHz
ReA post-accelerator
MHB
RB
*Production & In-flight separation
“Stopping” area
12 keV/u
12 keV/u x A
Charge breeding times: 200 – 20 ms
NSCL rare-isotope beam rates, < 109 pps
48Ca 0.3 - 20 MeV/u
238U 0.3 - 12 MeV/u
After production and separation, the beam is sent to the beam stopping area, decelerated, thermalized, and transported to the EBIT.

## Slide 7: The beam “stopping” area
DC beams > 80 MeV/u
He gas cell (Argonne lab)
Rare-isotope beams
from the production area
Analyzing dipole magnet
Si det. measure b activity for particle ID & beam transport optimization
DC beams < 60 keV
After thermalization, the beam is transported thru a low-energy beam line to a analyzing dipole for mass selection.
Beam thermalization:
Decelerate rare-isotope beams
Reduce emittance to 2p mm mrad (95%, 30 keV) before reacceleration
For efficient beam transport and EBIT injections
We have Si det. To measure the activity of the beam and optimization the beam transport and ID the type of beam we are dealing with.
The purpose of beam thermalization…

## Slide 8: The beam “stopping” area
Second beam line:
Equipped with offline (surface) ion source (e.g., K, Rb…)
For commissioning purposes
Analyzing dipole magnet
DC beams < 60 keV
On the other side of the beam lines…

## Slide 9: From the stopping area to ReA
Analyzing
magnet
To post-accelerator
built on an elevated platform
DC beams < 60 keV

## Slide 10: EBIT charge breeder in the ReA facility
Q/A separator
Room-temperature RFQ accelerator
SRF linear accelerator (LINAC)
Gas-cell stopper
EBIT charge breeder
Why do we use a charge breeder & why an EBIT (& not an Electron Cyclotron Resonance Ion Source (ECRIS) ) ?
  CB: Makes compact & cost-efficient compared with accelerator with stripper-foil stages.
  EBIT: High efficiency  Narrow charge state distribution (less ions lost in many charge states)
  EBIT: Fast  Breeding times < 50 ms (adjustable)
  EBIT: High beam purity / Low contamination level  e.g., cryogenic trap (4K)
Installed a few weeks ago!

## Slide 11: EBIT working principle
Trap electrodes
Magnetic field,
up to several Tesla
Electron
collector
Electron
gun cathode
Radial space-charge potential from the electron beam
Electron
beam
Axial potential well from the trap electrodes
Highly charged ions
Hair-like electron beam
F~100 mm
  What is an Electron-Beam Ion Trap and how does it work?
  Produces & traps highly charged ions with a high-current density electron beam
  3 Parts: e-gun, trap + magnet , e-collector
  Strong magnetic field: Electron-beam compression & Ionization by electron impact
  Trap electrodes Axial ion confinement
  Electron beam  Radial ion confinement
Quasi-continuous injection & accumulation (~100 ms)
A+
Pulsed extraction
AQ+
F~6 mm

## Slide 12: How do we inject & extract ions ?
  Typical width of an extracted ion pulse
  Typical pulse width (~10 ms) defined by the time taken by the trapped ions to freely exit the trap
  Instantaneous rate of each pulse is too high for many nuclear-physics experiments
  New extraction schemes being developed to spread in time pulses up to several ms.
Inj. & accu. (~100 ms)
Injec./Accu
Injec./Accu
Injec./Accu
Extract
Extract
Extract
  Extract (< 1 ms)
Time
Open the trap…
~ 10 ms
q+
q+
U
U
1+
1+
2+
2+
1+
1+
Ions axially trapped with potential well created by 2 barriers
Over-the-potential barrier injection
“Quasi-continuous”
Lower-the-barrier extraction
“Pulsed extraction”
Time sequence

## Slide 13: The ReA EBIT
Trap structure
* 1.2 m long
* 23 Ti electrodes
Main requirements of a CB for ReA:
Breeding time < 50 ms (for short-lived isotopes)
Efficiency in single charge states: 20% - 50 %
Charge capacity: up to 1010 positive charges
Low contamination level.
ReA EBIT key parameters:
High electron current: 1.4 A thru a 4T
E-beam energy < 30 keV (e.g., Ne-like U82+)
Current density (750mA/4T): ~500 A/cm2
Reduced contamination: 4-K trap structure
Unique feature for charge breeders: Viewports for visible, X-ray or gamma spectro. of highly charged radioactive ions.

## Slide 14: ReA EBIT: Double-magnet configuration
e-beam
i-beam
Low field
High field
Preferred B-field distribution
Over-the-potential barrier injection
U
U
1+
1+
2+
2+
1+
1+
Solenoid: Long low-field region
  To maximize ionization of 1+ ions within a roundtrip:
  Keep e-beam diameter for high electron-ion beam overlap upon injection  High capture probability
Helmholtz coils: Short high-field region
  Reduce e-beam diameter for high current density  High charge states and fast charge breeding
Monte-Carlo simulations of capture eff.:
E. Gavartin – M.Sc. Thesis, MSU,
K. Kittimanapun – PhD, MSU.
Capture efficiency
from Monte-Carlo simulations
@1A: >80% for <3 p mm mrad
Current density~1000 A/cm2
“Quasi-continuous”
Location of the two field regions
Notes: NSCL unique manget configuration
We inject ion over a trapping barrier, so to ionize them before a round trip, we an an extended low-field region.

## Slide 15: High-current electron gun to maximize the capture efficiency of injection ions
Demonstrated stable operation
up to ~ 1.4 A (4 T)
  Thermionic electron gun
  Ba-dispenser cathodes
  6-mm in diameter
  Modular cathode assembly
Perveance ~ 1.7 A/V3/2
  Anode
  Focus elec.
  Cathode
Modular ass’y includes:
E-beam current vs. extraction voltage
Can go away!

## Slide 16: Injection modes
Three modes of injection to produce stable-isotope and rare-isotope beams
  Gas injection producing an atomic beam
  Inject singly charged stable-isotope ions from two offline ion sources
  Inject rare-isotope beams from the beam stopping area
  Simultaneous injection of multiple beams from different sources…
Gas inlet
Offline ion source
Electron gun
Magnet cryostat
Offline source in the stopping area
Rare-isotope beams
Notes: E-beam character, Helmholtz coil compress e-beam, radially trap by neg space chage from e beam and radially by trap electrodes
High j rapidly increase charge states for short-lived isotopes <10ms, radial trap axial trap
MSU EBIT is different from mag field config and improved e-gun design. E-gun is mount on a plug-type connect, convenient for future upgrade
For fast charge breeding, long trap, mixing the idea of EBI s and t

## Slide 17: Charge-over-mass (Q/A) separator
Design parameters
  Measured resolving power, (A/Q)/D(A/Q)~500
  Achromatic within DE/E~3%
  Large energy acceptance to accept EBIT beams
  EBIT beam energy spread (~ion temperature), DE<500 eV  DE/E~0.3%
Dm
Double-focusing (achromatic) spectrometer
90o dipole magnet
45o electrostatic bends
Dm

## Slide 18: Many other EBIS/T around the world…
Brookhaven
RHIC-EBIS
TRIUMF/
Vancouver
LLNL
Livermore-EBIT
E-beam energy ~ 200 keV
Spectroscopic measurements in Bare Uranium
NSCL/FRIB
ANL
Applications
Atomic spectroscopy / Surface interaction studies
Charge breeder (source of stable & rare isotopes)
Charge breeder (post-acceleration of rare-isotope beams)
TITAN-EBIT
Tokyo
Geneva
GSI/Jena
Frankfurt
Shanghai
Dubna
Dresden
Heidelberg
Kielce
Clemson
NIST
LBL
REX-EBIS @ CERN
HD-EBIT
Flash-EBIT/SLAC
To mention: REX-EBIS@CERN: First EBIS breeder (Op. ~2000)
REA EBIT: Commissioning/entering operation
CARIBU, Argonne Nat. Lab.: Being commissioned
TRIUMF/ARIEL/CANREB (Planning)
KoRIA (Construction)
TRIUMF’s CANREB (2nd)
Only 5 breeders of RIB for post-acceleration:
Notes: EBIT/S are widely used around the word. Many of them are used in AMS. A few of them are used to study surface interaction and some are used as a charge breeder. More specifically, there are only 3 facilities around the world which are cern, our lab, and argonne using these device as a charge breeder for post accelerator for rare isotope beam. REXEBIS at cern successfully demonstrated EBIS CB for rib as a first group.

## Slide 19: “Charge Breeding” Physics Quick overview of what happens during charge breeding

## Slide 20: Charge breeding physics
Charge breeding in the EBIT is mainly governed by 4 processes
  Electron-impact ionization
  Radiative recombination with beam electrons
  Charge exchange (elec. capture) with residual gas (neutral)
  4. Ion escape (evaporation) from the trap
Charge state evolution described by coupled differential rate equations
e-impact ionization
radiative recombination
charge ex. recombination
Ion “evaporation”
E-beam current density
Residual gas pressure
Trapping potential: Can bee minimized for high V trap
AQ+
AQ+1
AQ-1
Ionization
Ionization
Recombination
Recombination
Ionization
Ionization
“Evaporation”
“Evaporation”
“Evaporation”
Each ion species has a its own rate equation
In this equation that describe time variation of the population of an ion species
E-impact and rr depends on the density
Each species has its own rate equation
To a good approximation this contribution can be neglected!

## Slide 21: How crucial is each atomic-physics process?
Depends on the cross sections, which control the production rates of every charge state
Radiative recombination:
CS small for low charge states
(contribution can be neglected for simplicity)
Electron-impact ionization
(Ee=30 keV ~ 1.7 IP)
Charge exchange with residual gas:
High CS but Loss rates small for low pressure
(contribution can be neglected for simplicity)
  To a good approximation: charge breeding dominated by e-impact ionization for high vacuum & low charge states
Krypton

## Slide 22: Why is the e-impact ionization cross section decreasing for high charge states ?
Electron-impact ionization cross-section:
Ee: E-beam energy
IP: Ionization potential
Ionization threshold
2.7x
The cross section depends on the ion radius:
*E-impact cross section is function of the ionization potential, increasing for high charge states*
n=4
Ni-like Kr8+
Ne-like Kr26+
He-like Kr34+
n=1
n=2
n=3
Need lots of energy to ionize highly charged ions
~20 keV is needed to break the n=1 elec. shell
Krypton
Another way to seen it:

## Slide 23: Charge breeding times
0
1
Time for full conversion (= assuming
Charge breeding times are inversely proportional to the e-beam current density (je), and e-impact ionization cross sections (sEI) (neglect all other processes)
Ionization cross sections are fixed atomic parameters…
The current density is a device parameter that can be optimized…
R&D program to develop e-guns of high-current & high-current density (~104 A/cm2)
How can we reduce the CB times ?
ReA EBIT, Average ~500 A/cm2
~100 ms
~2 s
~40 s
Charge breeding for reacceleation mostly limited to moderate charge states: e.g, Kr26+ , K17+ , etc…

## Slide 24
Physics of Continuous Beam Injection

## Slide 25: A geometrical model for the capture efficiency of continuous beam injection
EBIT acceptance:
 Describes e-beam & injected ion beam overlap
Geometrical eff.:
Ionization probability:
Need high current, Ie & large e-beam radius, re
Need high density, je
(imply a small e-radius)
U
1+
troundtrip ~ 50 mm
The acceptance and the ionization probability have opposite requirements!
To reach high capture efficiencies, need both: high e-beam current AND high-current density!
(5% instead of 10% in efficiency, make a difference between one or two days of data taking of the users)
Beam emittance: Area in phase-space of the ensemble of beam particles
x
vx
Acceptance: Phase-space area of an ensemble of particles a device can accept
Capture
EBIT beam acceptance: F. Wenander et al. , REX-EBIS, Design & Simulations, CERN 1999
Total capture efficiency:

## Slide 26: Simplified model…
  Typical ion beam emittance < 3 p mm mrad (95%, 30 keV)
  Assuming a perfect overlap between the ion and electron beams
  Acceptance > Emittance  Geometrical Eff. ~1
Where is the ReA EBIT ???
… in the next slides…
  Total efficiency vs. e-beam current density
x
vx
Total Eff. ~ Ionization probability

## Slide 27: Commissioning results

## Slide 28: Three research areas we focus on to improve performance…
Charge Breeding Efficiency (Capture & Breeding Time)
Beam Purity
(No Contamination !)
Optimize the Width of the Extracted Ion Pulses

## Slide 29: Measured charge breeding efficiency
With careful optimization:
Up to 6% for K15+
Capture eff.: ~30%
Capture efficiency:
Field config.: 4T (H-coils)–2T (sol.)
Electron-beam current: 750 mA
Electron-beam energy: 15.5 keV
Accumulation/Breeding time ~100 ms
K15+
Efficiency in single charge states of injected 39K stable-isotope beams
   Measured with a Faraday cup after Q/A separator
Does it make sense ?

## Slide 30: Electron-beam radius measurement and current density
X-ray image of the electron beam
  The capture efficiency of injected beams varies with the e-beam current density
  The density determined from the known current & e-beam radius 
  The e-beam radius measured by imaging x-rays emitted by trapped highly charged ions.
Electron beam
Slit
(Spatial filter)
Projected x-ray
intensity profile
CCD chip on
the image plane
X-ray pinhole camera set-up on the ReA EBIT
X-ray camera
EBIT
X-ray image of the electron beam
  je ~ 454(83) A/cm2 (4 T)
  re (80%) ~ 212(19) mm for 800 mA, 4-T field (Helmholtz)
  T. Baumann, A. Lapierre, S. Schwarz, K. Kittimanapun, D. Leitner, G. Bollen, RSI 85, 073302 (2014)

## Slide 31: Calculated capture efficiency of continuously injected beams
  Av. je ~ 320 A/cm2 for 2T (sol.) - 4 T(Helmholtz) field config.
  re~212 mm  Acceptance ~15 p mm mrad
  Acceptance > Emittance, Geometrical Eff. ~1
U
1+
troundtrip ~ 41 mm (K+)
Total Eff. ~ Capture Eff. by ionization
For ~320 A/cm2, we can expect at most ~30%
Agree with expt’l results!!!

## Slide 32: Calculated capture efficiency of continuously injected beams
  Av. je ~ 320 A/cm2 for 2T (sol.) - 4 T(Helmholtz) field config.
  re~212 mm  Acceptance ~15 p mm mrad
  Acceptance > Emittance, Geometrical Eff. ~1
U
1+
troundtrip ~ 41 mm (K+)
Total Eff. ~ Capture Eff. by ionization
For ~320 A/cm2, we can expect at most ~30%
Agree with expt’l results!!!
How can we get to 80% ?

## Slide 33: Higher capture eff.: Option 1  Higher current density
3 basic strategies to increase the capture efficiency of beam continuously injected beams
Compress electron-beam radius
Increase electron-beam current
Both: High compression (6T) & current (1.4A)  Will be investigated in the next months.
Acceptance > Emittance
Fixed trap length: ~ 644 mm
Calculated capture efficiency vs. e-beam current & radius
Eff. 30%: 330 A/cm2
Eff. 80%: 1545 A/cm2

## Slide 34: Higher capture eff.: Option 2  Inject ion pulses
Beam line for continuous injection
Analyzing
magnet
To ReA
DC beams < 60 keV
To EBIT
Installing Radio-frequency quadrupole (RFQ) “cooler-buncher” ion trap:
Accumulate injected beams
Produce ion pulses with a cooling gas (e.g., He)
Under installation / Start commissioning ~ Mid-Dec
Engineering model

## Slide 35: Higher capture eff.: Option 2  Inject ion pulses
Q+
Q+
U
U
1+
1+
2+
Q+
Pulsed injection scheme
Charge breeding
Pulsed extraction
1+
1+
Pulsed injection
Expected capture eff. of injected beams: >90%
Expected breeding eff. in single charge states: up to ~20%
Max. ~105 - 108 ions/pulse
Max. ~ 107 - 1010 ions/sec @100 Hz (~max. rate)
Not efficient for high injected beam rates.
However, the “buncher” ion trap has a space-charge limit…
***Continuous beam injection is the best method for high beam rates expected with FRIB***
***Still need to optimize continuous ion beam injection***

## Slide 36: Time stretching of EBIT extracted pulses
Inj. & Accu. (~100 ms)
  Extract (< 1 ms)
Time
Slow ion release over ~ ms
Ramp-like pulse (AFG)
~ ms
Train of pulses with a series of successive ~ms square pulses
  Conventional
  Slow ramp
  Train (or Comb)
Time
Trap opened with a square pulse
  Developing schemes to spreading in time extracted ion pulses is important for efficiency detection by user’s detection systems:
Reduce the instantaneous rate (# of ions per pulse (time) width)  Conventional pulse ~10 ms
DAQ systems can have long dead times (~10 ms) preventing detection of all ions within a pulse.
Time

## Slide 37: Time stretching of EBIT extracted pulses
Width ~ 10 ms
Each pulse is ~2-ms long and contains only a few ions
T. Baumann, A. Lapierre, S. Schwarz, et al. AIP Conf. Proceedings (EBIST2014 International Symposium), to be published (2015)
Measured with a microchannel plate after the Q/A separator
Peak intensities normalized to 1
Each pulses spaced by 100 ms (variable)

## Slide 38: Contamination maps
Q/A
separator
Measure beam energy after RFQ (ERFQ) with Si detector
EBIT
RFQ
Identifying contaminants is essential to eliminate them & provide rare-isotope beams of high purity
RFQ accel. energy:
Separator measures the Q/A:
40Ar13+
40Ar14+
12C4+
16O5+
16O6+
63Cu20+
38Ar13+
35Cl12+
32S10+
23Na8+
No trace of Co, Ti, Ca here!
28Si9+
19F6+
Krypton isotopes from injected gas
Contaminants:
Na, Cl  NaCl (fingerprint?)
Si  Aluminized mylar tape adhesive
F  Scroll-pumps: Bearing lubricant & tip seal (Teflon)
S  Dichronite UHV bearing lubricant (WS2)
Ba, W  Dispenser cathode
All C, N, O, Ar stable isotopes  form residual gas.
Program to map Q/A regions

## Slide 39: Charge breeding of rare-isotope beams
April 2013: Charge breeding of 76Ga [t1/2=33 sec]
Q/A separator
RFQ
SRF cryomodules
Gas-cell stopper
EBIT
b-decay energy spectra
Charge state distribution after Q/A separator
b-activity measured
with decay counter
Notes: In these two figures we can see
In the right figure we have a CSD distibution …, where we can see the charge state as a function of the beta activity of the beam

## Slide 40: Delivery of rare-isotope beams
August 2013: First beam delivery to users for studies of nuclear reactions
ANASEN: Array for Nuclear Astrophysics Studies with Exotic Nuclei
ANASEN
Charge state distribution after Q/A separator
Charge breeding and reacceleration of 37K17+ [t1/2 ~1 s] to 2.3 MeV/u
Total E (chan)
ΔE (chan)
37K17+
ANASEN gas chamber (prop. counter)
37Cl17+
13C6+
EBIT
b-activity measured
with decay counter
Rate~1000 pps

## Slide 41: Summary & Outlook
  EBIT being commissioned
  ReA post-accelerator can deliver rare-isotope beams…
  Superconducting magnet failed in Jan. 2014, being repaired, expect return in Feb. 2015
  Next year, EBIT work focus on:
    Improving the charge breeding efficiencies (continuous & pulsed injection)
    Stretching the extracted pulse structure
    Improving the beam purity
    Delivering beams to users
  Early science program with rare-isotope beams next year…
Thank you.

## Slide 42: Back-up slides

## Slide 43: Reacceleration of stable-isotope beams
Dec. 2012: Demonstrated charge breeding of Rb+ stable-isotope ions from the offline source in the beam stopping area.
April 2013: Reacceleration of 85Rb27+ & 87Rb28+ with the RFQ & SRF cavities.
***First time the beam stopping area was coupled to ReA as a single unit***
Q/A spectrum of charge bred Rb
Energy spectrum of reaccelerated 87Rb28+ with Si detector after LINAC
Notes: The left figure shows…
Produced from beams of residual gas in the EBIT: O and Ar
Energy spectrum of reaccelerated 87Rb28+ taken with a si detector after reacceleration…

## Slide 44: Calculated capture efficiency
  From electron-beam size measurements, average current density, J~400 A/cm2
  For beam acceptance > injection beam emittance (ion-electron overlap), Geometrical Eff. ~1
trf: Interaction time of 1+ ions with e-beam before they can exit the trap.
U
U
1+
1+
This curve is telling us that for 400 A/ cm2, we can only expect at most ~30%.
trt~41 mm (K+)
We are here !!!
Total Eff. ~ Capture Eff. by ionization = 1 - exp[-s1-2 trt J/e]
How can we get to 80%, at least ?

## Slide 45: Reacceleration of rare-isotope beams
April 2013: Reacceleration of 76Ga24+ & 76Ga25+ to 1 MeV/u
Q/A separator
RFQ
SRF cryomodules
Gas-cell stopper
EBIT
b-decay counter
b-decay energy spectrum after reacceleration

## Slide 46: NSCL coupled SC cyclotrons

## Slide 47: Capture efficiency (continuous injection) vs length of the high-field region
Chart: Ratio of the Helmholtz coils (high-field) to the solenoid (low-field) lengths [%]
| 1 | 5 | 10 | 15 | 20 | 25 | 30 | 40 | 60 | 80 | 99
6T-0.5T/HC/250mA/1m | 56 | 68 | 77 | 81 | 82 | 81 | 80 | 75 | 64 | 52 | 40
6T-2T/HC/250mA/1m | 66 | 66 | 65 | 64 | 62 | 61 | 60 | 57 | 51 | 45 | 40
6T-4T/HC/250mA/1m | 48 | 48 | 48 | 47 | 47 | 46 | 46 | 45 | 43 | 41 | 40
6T-0.5T/HC/500mA/1m | 75 | 87 | 91 | 92 | 91 | 89 | 87 | 82 | 74 | 65 | 56
6T-2T/HC/500mA/1m | 98 | 97 | 94 | 92 | 90 | 88 | 86 | 82 | 73 | 64 | 56
6T-4T/HC/500mA/1m | 68.52 | 68.01 | 67.37 | 66.74 | 66.1 | 65.46 | 64.84 | 63.55 | 61 | 58.45 | 56.03
6T-0.5T/HC/1A/1m | 88 | 96 | 98 | 97 | 96 | 95 | 94 | 92 | 88 | 84 | 80
6T-4T/HC/1A/1m | 100 | 99 | 98 | 97 | 96 | 95 | 94 | 92 | 88 | 84 | 80
6T-2T/HC/2A/1m | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100
4T/LC/1.2A/1m | 79
2T/LC/2.4A/1m | 78
6T/HC/2A/0.1m | 100
Full Helmholtz
(High field)
Full Solenoid
(Low field)

## Slide 48: CBSIM – Charge Breeding Simulation for Kr

## Slide 49
Ionization cross sections depend on beam energy (Ee), ionization potential (Ij), and number of bound electrons (rj) in each atomic shell j
[cm2, eV]
Ionization cross section  1/q2 → rapid decrease as charge state increases
Breeding high charge states by successive ionization
Electron energy (keV) | Ionization potential (keV) | Cross section (cm2) | Rate per ion (s-1) at 5000 A/cm2
10 | 0.02 | ~ 10-17 | ~ 105
200 | 130 | ~ 10-24 | ~ 10-2
Time for sequential ionization to charge state n
Example: Sn(Z=50) in a 2000 A/cm2 beam
 150 ms to reach Ne-like Sn40+

## Slide 50: Contamination measurements
  Rare isotopes injected at low rates (< 106 ions/sec)
  Delivering clean beams to users is important
  Program: Quantify Q/A regions the least contaminated w/ single-ion counting techniques
Q/A scans with a high-sensitivity ammeter
37K17+ : 1000 pps.
Request < 1% of 37Cl (not met)
Request < 50% for low-Z ions
To give an example how critical it is to deliver clean beam…
Many Q/A regions with a level of contamination of < 0.1 pA
Notes: In this figure we can see peaks of the most abundant (well-known) stable isotopes of the residual gas in the ebit, but also not so adundant isotopes…
What we can say from these figures is that the level of contamination of many region is less than 0.1 pA, but than are many contaminants than can only be meaured with single-ion techniques

## Slide 51: CBSIM – Charge Breeding Simulation for Kr
He-like
Ne-like
0
1
2
3
4
5
6
26
34
Ee=30 keV
Je =500 A/cm2
Typical charge breeding times:
Ne-like Kr26+ ~100 ms
He-like Kr34+ ~2s
Bare Kr36+ ~40s

## Slide 52: Commissioning results
K18+ beam on MCP viewer
Emittance vs charge state
Transverse emittance measurements with Pepperpot-meter
Preliminary
Notes: E-beam character, Helmholtz coil compress e-beam, radially trap by neg space chage from e beam and radially by trap electrodes
High j rapidly increase charge states for short-lived isotopes <10ms, radial trap axial trap
MSU EBIT is different from mag field config and improved e-gun design. E-gun is mount on a plug-type connect, convenient for future upgrade
For fast charge breeding, long trap, mixing the idea of EBI s and t

## Slide 53: The EBIT electron current density is key for optimizing the charge breeder efficiency for ReA
  Breeding times and capture efficiencies are directly related to the current density, which is a function of the electron-beam current and radius.
    The higher the electron-beam current density the shorter the charge breeding times and the higher the capture efficiencies of ions injected into the ReA EBIT charge breeder
Simulated charge-breeding time of Krypton
where
Calculated capture efficiency of injected 1+ ions
A. Lapierre and S. Schwarz (ReA EBIT group), ReA Department

## Slide 54: A feel about the breeding time…
M. W. Froese, M Master Thesis, University of Manitoba (based on R. Becker’s CBSIM program)
Ionization time to reach the population maximum
Note: Recombination with beam electrons and charge exchange with atoms and molecules not included
He-like Kr ~30 ms

## Slide 55: Capture efficiency vs. Electron-beam energy
R224
4T 0.8 501
1.28 m round trip
3 pi mm mrad
Chart: Electron beam energy (keV)
| 0.5 | 1 | 4 | 8 | 12 | 16 | 20
0.3 A | 88 | 72 | 36 | 23 | 18 | 14 | 12
0.6 A | 99 | 92 | 59 | 41 | 32 | 27 | 23
0.8 A | 100 | 97 | 69 | 50 | 40 | 33 | 29
1.2 A | 100 | 99 | 83 | 65 | 54 | 46 | 41

## Slide 56: Delivery of rare-isotope beams
Total beam energy vs. energy loss in the ANASEN gas chamber (prop. counter)
ΔE (chan)
Total E (chan)
37Cl17+
13C6+
No K beam
Total E (chan)
ΔE (chan)
37K17+
With K beam
15N7+ ???

## Slide 57: CBSIM – Charge Breeding Simulation for Kr
Ions distributed over a multitude of charge states (not very good…intrinsic efficiency loss)
Injection
Similar graph, but we can see here how the charge state distribution evolves with time…
Log(t/10 [sec])

## Slide 58: Capture efficiency of continuously injected beams
Width: Acceptance (Pgeo)
Height: Ionization probability (P12)
All EBIT breeders operating in the low-density regime  Preferred for high acceptance !
Real reason!
 Injecting high-current e-beams (~ 1 A) into strong magnetic fields is challenging.
Current densities of ~104 A/cm2 has never been reached worldwide.
Small re / High current density:
High eff. for small emittances, only!
(& perfectly aligned beams)
Large re / Low current density:
Higher eff. for large emittances
(…more alignment forgiving!)
Total efficiency vs. ion beam emittance
You don’t to remember all the equation, but you have to remember that if the electron beam is large , total = capr by ionzation
ReA EBIT
Can go away!

## Slide 59: Expected FRIB rare-isotope yields
Rates are available at http://groups.nscl.msu.edu/frib/rates/
Current NSCL rates ≤ 109 pps
FRIB will allow us to increase the current production rates by up to two orders of magnitude
Slide can go away!

## Slide 60
What do highly charged ions look like?
Examples of one-electron atoms
Highly charged ions:
  Very high electron binding and transition energies
  Very small ionization cross sections
  Much smaller than neutral atoms
Courtesy of R. Marrs, co-inventor of the Livermore-EBIT at the Lawrence Livermore National Lab.
rH / rU 91+ ~ 92
RSun / REarth ~ 109
Earth
The relative size of the H atoms wrt to H-like U, which has a ratio in radius of 92 and very close to the ratio of sun radius to the earth radius.
This slide also the ratio is ionization cross section which is 10000x less for H-like U91+

## Slide 61: CBSIM – Charge Breeding Simulation for Kr
Log(t/10 [sec])
He-like Kr34+
Bare
Ne-like Kr26+
Charge breeding times:
Ne-like Kr26+ ~100 ms
He-like Kr34+ ~2s
Bare Kr36+ ~40s
Injection
We can simulate the CB process
Ion population maximum of Ne-like Kr: trap contains a narrow charge state distribution  intrinsic loss mechanism