[image: image10.png]



GRETINA Digitizer Specification

Document # GRT-

-PRELIMINARY-

August 15, 2006

Dionisio Doering, John Joseph, Harold Yaver, Sergio Zimmermann

Accepted and Approved by:

Sergio Zimmermann __________________________Date________________

I-Yang Lee  _________________________________Date_________________

GRETINA Digitizer Specification

Gamma Ray Energy Tracking In-Beam Nuclear Array (GRETINA)

Revision History

	Revision No.
	Pages Affected
	Effective Date

	
	
	

	
	
	

	
	
	

	
	
	

	
	
	

	
	
	

	
	
	

	
	
	

	
	
	

	
	
	

	
	
	

	
	
	

	
	
	

	
	
	

	
	
	

	
	
	

	
	
	



51
INTRODUCTION


51.1
OVERALL DESCRIPTION


51.2
PRECEDENCE


52
DIGITIZER MODULES


52.1
GRETINA DIGITIZER TOP LEVEL BLOCK DIAGRAM


62.2
FUNCTIONAL BLOCKs


62.2.1
VMEBUS INTERFACE


72.2.1.1
VME INTERFACE DESCRIPTION


82.2.1.2
VME TO MAIN FPGA TIMING


92.2.1.3
GRETINA Digitizer Memory Map


92.2.1.4
VME FPGA Memory Map


112.2.1.5
VMEBUS J1/P1 CONNECTOR PIN ASSIGNMENTS


122.2.1.6
VMEBUS J2/P2 CONNECTOR PIN ASSIGNMENTS


132.2.2
FRONT END PREAMP AND DIGITIZER


132.2.2.1
ANALOG INPUT


152.2.2.2
FRONT END PREAMP


162.2.2.3
CENTRAL CONTACT REPEATER OUTPUT


162.2.2.4
DIGITIZER CIRCUIT


162.2.3
EVENT FIFO


162.2.4
MAIN FPGA


192.2.4.1
MODES OF OPERATION


192.2.4.1.1
INTERNAL MODE


192.2.4.1.2
EXTERNAL MODE


202.2.4.1.3
INTERNAL WITH VALIDATION MODE


202.2.4.1.4
TTCL MODE


212.2.4.2
CHECKING SYNCHRONISM


212.2.4.3
INITIALIZATION


212.2.4.3.1
SYSTEM INITIALIZATION


212.2.4.3.2
INITIALIZATION OF TTCL COMMUNICATION


222.2.4.4
CLOCK MANAGEMENT


222.2.4.5
OUTPUT DATA FORMAT


242.2.4.6
MAIN FPGA REGISTERS


282.2.5
DAC OUTPUTS


292.2.6
FRONTBUS INTERFACE


332.2.6.1
FRONT BUS REGISTERS


332.2.7
SERIALIZER/DESERIALIZER


342.2.8
AUXILLIARY INPUT INTERFACE


342.2.9
100MHz CLOCK DISTRIBUTION


342.2.10
FRONT PANEL LED INDICATORS


352.2.11
VOLTAGE MONITORS


352.2.12
TEMPERATURE MONITORS


353
REFERENCES




INTRODUCTION

1.1 OVERALL DESCRIPTION

The GRETINA Digitizer module is a combination of digitizer and digital signal processor.  It accepts 10 inputs directly from the detector module pre-amplifiers and digitizes at a nominal frequency of 100 MHz with 14 bits ADC precision.  The ADC counts interface to the FPGA, which digitally processes the data.  

1.2 PRECEDENCE

The design of the GRETINA Digitizer electronics board must meet the Digitizer specific requirements defined in the GRETINA Electronics Requirements Document.  

2 DIGITIZER MODULES

2.1 GRETINA DIGITIZER TOP LEVEL BLOCK DIAGRAM


[image: image1.wmf]GRETINA DIGITIZER BLOCK DIAGRAM

CH0 FE/ADC

DAC0 OUT

LED OUT

AUX IO IN

FRONTBUS  IO

SER/DES   IO

CC OUT

EVENT 

FIFO

50MHz REF CLK

50MHz RCLK

CH1 FE/ADC

CH9 FE/ADC

DAC1 OUT

VME 

CTRL

& 

I/O

MAIN FPGA

XC3S5000

-

5FG900

(633 USER I/O)

100MHz CLOCK DIST

MUX


Figure 1: GRETINA DIGITIZER BLOCK DIAGRAM

2.2 FUNCTIONAL BLOCKs

The LBNL GRETINA Digitizer Electronics board will include all of the functions and blocks described in the following sections and listed below.

VMEBUS INTERFACE
Interface to the Host Computer System

FRONT END PREAMP AND DIGITIZER
Interface to the FE Pre Amplifiers of the Crystal Detectors

EVENT FIFO
256K x 32 FIFO used to buffer Event Data while the VME Host is fetching hit data

MAIN FPGA
Real-Time FPGA Signal Processor

DAC OUTPUTS
Geneal Purpose Current or Voltage output

FRONTBUS INTERFACE
Interface used by all Digitizers in one crate to communicate and share vital information

SERIALIZER/DESERIALIZER
Interface to the Trigger and Timing System

AUXILLIARY INPUT INTERFACE
16 General Purpose TTL Inputs to the MAIN FPGA

100MHz CLOCK DISTRIBUTION
Onboard Clock Management

FRONT PANEL LED INDICATORS
User defined Front Panel Monitoring

VOLTAGE MONITORS
Monitoring of voltage level of all key power supply voltages

TEMPERATURE MONITORS
Simple over-temperature switches

2.2.1 VMEBUS INTERFACE

The GRETINA Digitizer board will communicate with a host computer using the VME64x protocol.  The slave interface will support the A32/D32 address and data format for block transfers and single read/write access.  Geographical addressing (GADDR) will be used to determine the base address of each board in a crate, so any Digitizer can be used in any slot without the need to set a board address on the electronics card.  The block diagram of the VME Interface is shown in Figure 2.


[image: image2.wmf]GRETINA VME INTERFACE BLOCK DIAGRAM

MAIN FPGA 

CONFIG. 

MEMORY 

(128/64/32

Mbits

)

VME 

BUS 

I/O

VME FPGA

XC3S400

-

4FG320

(221 USER I/O)

50MHz 

CLOCK 

DIST

16

DATA

CTRL

BUSY

MAIN FPGA CONFIG

4

5

32

DATA

32

ADDR

6

AM

VME CTRL

7

50MHz REFERENCE CLOCK

DATA

ADDR

32

32

GADDR

6

VME SPEC

8

MEMORY MAP CTRL

3

SYSTEM MONITORING

6

IRQ

10


Figure 2: GRETINA VME INTERFACE BLOCK DIAGRAM
2.2.1.1 VME INTERFACE DESCRIPTION

The design of the VME interface on the Gretina Digitizer will be implemented using a Xilinx Spartan3 FPGA and a small number of buffers, drivers and transceivers that are required to physically interface to the VME back plane.  The VME FPGA will control the block transfer and programmed IO read/write access to the Digitizer Main FPGA, the configuration Flash memory, and the VME FPGA internal register memory map.  The VME FPGA will include functional blocks that will control reset and configuration of the Main FPGA at system power-on and on command when requested by the VME host.  The 3 main addressable locations on the Digitizer board are the Main FPGA, the VME Controller FPGA, and the Main FPGA Configuration Memory.  Control of the on-board peripherals will be provided through the Main FPGA memory map.  These peripherals include the Front End ADC circuits, the output DACs, the FrontBus interface, and the Main Readout FIFO to list a few.

The VME interface requires a stand-alone 50MHz clock that will be buffered through a 1 to 4 low skew differential clock distribution circuit.  The 50MHz clock signal will be routed to the Main FPGA and to the Serializer/Deserializer (SER/DES) block to satisfy the operational requirements of each sub-system.   A constant clock is required by the Main FPGA to allow readout of internal status registers in the case that the recovered SER/DES clock falls out of lock or is not present. 

The configuration memory will be implemented with one Intel 128/64/32Mbit, 64k sector flash memory component and will be used primarily to hold configuration programs for the Main FPGA.  The Main FPGA Configuration Flash will be accessible on the VME bus for Read/Write operations as required.

2.2.1.2 VME TO MAIN FPGA TIMING

The following timing diagrams show representations of programmed I/O (single data transfer) access only.  Block transfer is supported for all registers in the MAIN and VME FPGA.  The Block Transfer timing for transfers from the MAIN FPGA to the VME FPGA is not shown because the VME FPGA will handles all the control protocol as long as the MAIN FPGA asserts LACK for every valid DATA STROBE/CYCLE.  Block Transfer will be allowed for read access to the MAIN FPGA configuration Flash memory.  Block Transfer for write access to the MAIN FPGA configuration Flash memory may not be supported.  

Local VME Read Cycle

                           _________________________________________________

  VME_ADDR(31:1) XXXXXXXXXX_________________________________________________XXXXX

                 ________________________________________________________________

  VME_RNW        

                 _____________________                                    _______

  VME_DBE_N(3:0)                      \__________________________________/

                 ___________________________________________________          ___

  VME_DTACK_N                                                       \________/

                                                             ___________________ 

  VME_DATA(31:0) -------------------------------------------<___________________>-

The VME FPGA generates VME_DTACK_N for completed valid cycle.

VME FPGA to MAIN FPGA Read Timing

                        _______________________________________________

  LADDR(31:0) XXXXXXXXXX_______________________________________________XXXXXXXXXXX

                            _________________________________________

  CS(0)       _____________/                                         \____________

              ____________________________________________________________________

  RNW        

                                         _________________   

  STRB        __________________________/                 \_______________________

                                                    __________

  LACK        _____________________________________/          \___________________

                                                ___________________ 

  LDATA(31:0) ---------------------------------<___________________>--------------

                __    __    __    __    __    __    __    __    __    __    __    

  CLK50      __/  \__/  \__/  \__/  \__/  \__/  \__/  \__/  \__/  \__/  \__/  \___

The MAIN FPGA generates LACK when LDATA is valid on the data bus.

Local VME Write Cycle

                    _____________________________________________

  VME_ADDR(31:1) XXX_____________________________________________XXXXXXX

                 ______                                          ______

  VME_RNW              \________________________________________/

                 ____________                                  ________

  VME_DBE_N(3:0)             \________________________________/

                 __________________________________________       _____

  VME_DTACK_N                                              \_____/

                          ________

  VME_DATA(31:0) --------<________>------------------------------------

The VME FPGA generates VME_DTACK_N for completed valid cycle.

VME to MAIN Write Timing

                          ______________________________________

  LADDR(31:0)   XXXXXXXXXX______________________________________XXXXXXXXXXX

                                  _____________________________

  CS(0)      ____________________/                             \___________

             ______________                                           _____

  RNW                      \_________________________________________/

                                              ________   

  STRB       ________________________________/        \____________________

                                                    ____

  LACK       ______________________________________/    \__________________

                                        ___________________ 

  LDATA(31:0)   -----------------------<___________________>---------------

                __    __    __    __    __    __    __    __    __    __       
  CLK50      __/  \__/  \__/  \__/  \__/  \__/  \__/  \__/  \__/  \__/  \__

The MAIN FPGA generates LACK when LDATA is latched in the device.

2.2.1.3 GRETINA Digitizer Memory Map

The Geographical Address bits on the VME back plane determine the Base Address of the Digitizer boards.  The GA bits are compared to bit 31 to 27 of the VME address word.

Table 1 – Gretina Digitizer Memory Map

	GRETINA Digitizer: Top Level Memory Map Bit Definitions

	31:27
	26
	25
	24
	23:0
	Start Address/Description

	GA
	0
	0
	0
	Addressable Memory
	Base Address + 0x0000000: Main FPGA Registers

	GA
	0
	0
	1
	Addressable Memory
	Base Address + 0x1000000: VME FPGA Registers

	GA
	1
	1
	1
	FLASH ADDR
	Base Address + 0x7000000: Up to 16Mbyte Configuration Flash


2.2.1.4 VME FPGA Memory Map

	Description
	Address
	Access
	Width

	DIGITIZER MAIN FPGA Configuration Control Register
	BASE_ADDRESS + 0x1000000
	RW
	32

	All bits in the Register are Self-Clearing, unless noted otherwise.
	Bit Value

	
	1
	0

	  Bit 0: Configure Main Digitizer FPGA

If this bit is set, the Main Digtizer FPGA will be erased and programmed with a Configuration Bit Stream that is stored in the on board flash memory
	Configure

Main FPGA
	Idle

	  Bit 1: Reset Main Digitizer FPGA

If this bit is set, a reset pulse is transmitted to the Main Digitizer FPGA
	Reset

Main FPGA
	Idle

	  Bits[31:2]: TBD
	
	


	Description
	Address
	Access
	Width

	DIGITIZER MAIN FPGA Configuration Status Register
	BASE_ADDRESS + 0x1000004
	R
	32

	Status of the MAIN FPGA Configuration
	Bit Value

	
	1
	0

	  Bit 0: MAIN FPGA DONE Bit Status
	DONE
	CNFG

	  Bit 1: MAIN FPGA INIT_N Bit Status
	ERASE
	OK

	  Bit 2: MAIN FPGA PROG_N Bit Status
	CNFG
	OK

	  Bit 3: MAIN FPGA HALT Config Status
	HALT
	OK

	  Bit 5: MAIN FPGA RESET Status
	RESET
	OK

	  Bits[31:4]: TBD 
	
	


	Description
	Address
	Access
	Width

	DIGITIZER Voltage and Temperature Status Register
	BASE_ADDRESS + 0x1000008
	R
	32

	Status of the Digitizer Voltage and Temperature
	Bit Value

	
	1
	0

	  Bit 0: All Power Supplies OK
	OK
	FAULT

	  Bit 1: Over Voltage Status
	FAULT
	OK

	  Bit 2: Under Voltage Status
	FAULT
	OK

	  Bit 3: Temperature Sensor 0 Status
	FAULT
	OK

	  Bit 4: Temperature Sensor 1 Status
	FAULT
	OK

	  Bit 5: Temperature Sensor 2 Status
	FAULT
	OK

	  Bits[31:6]: Not Used
	
	


	Description
	Address
	Access
	Width

	VME GP CONTROL Register
	BASE_ADDRESS + 0x1000010
	RW
	32

	General purpose VME Control setting
	Bit Value

	
	1
	0

	  Bits [31:0]: TBD
	VALUE


	Description
	Address
	Access
	Width

	VME IRQ CONTROL Register
	BASE_ADDRESS + 0x1000014
	RW
	32

	Interrupt setup for the Digitizer VME interface
	Bit Value

	
	1
	0

	  Bits [2:0]: IRQ value
	VALUE

	  Bits [31:3]: TBD
	VALUE


	Description
	Address
	Access
	Width

	FLASH MEMORY CONTROL REGISTERS
	BASE_ADDRESS + 0x1000020

to

0x100002C
	RW
	32

	Setup for the Digitizer Flash Memory Interface
	Bit Value

	
	1
	0

	  Bits [31:0]: TBD
	VALUE


	Description
	Address
	Access
	Width

	VME FPGA SAND BOX REGISTER BlOCK
	BASE_ADDRESS + 0x1000030

to

0x100003C
	RW
	32

	3 Registers to use for Test and Debug
	Bit Value

	
	1
	0

	  Bits [31: 0]: Registers for general purpose Read/Write access
	VALUE


	Description
	Address
	Access
	Width

	VME FPGA VERSION/REVISION Status Register
	BASE_ADDRESS + 0x100007C
	R
	32

	Status of the MAIN FPGA Configuration
	Bit Value

	
	1
	0

	  Bits [15: 0]: VHDL Version Number
	VALUE

	  Bits [23:16]: BOARD Revision Number
	VALUE

	  Bits [31:24]: TBD
	VALUE


2.2.1.5 VMEBUS J1/P1 CONNECTOR PIN ASSIGNMENTS

A 160 pin connector will be used for P1/J1 to allow use of the Geographical addressing pins and the 3.3V power supply available in the VME64x crates.

Table 2 – VME P1 Signals

	 VME I/O: J1/P1 Pin Assignments (VME64x)

	Pin
	Row Z
	Row A
	Row B
	Row C
	Row D

	1
	MPR
	D00
	BBSYn
	D08
	VPC

	2
	GND
	D01
	BCLRn
	D09
	GND

	3
	MCLK
	D02
	ACFAILn
	D10
	+V1

	4
	GND
	D03
	BG0INn
	D11
	+V2

	5
	MSD
	D04
	BG0OUTn
	D12
	RsvU

	6
	GND
	D05
	BG1INn
	D13
	-V1

	7
	MMD
	D06
	BG1OUTn
	D14
	-V2

	8
	GND
	D07
	BG2INn
	D15
	RsvU

	9
	MCTL
	GND
	BG2OUTn
	GND
	GAPn

	10
	GND
	SYSCLK
	BG3INn
	SYSFAILn
	GA0n

	11
	RESPn
	GND
	BG3OUTn
	BERRn
	GA1n

	12
	GND
	DS1n
	BR0n
	SYSRESETn
	+3.3V

	13
	RsvBus
	DS0n
	BR1n
	LWORDn
	GA2n

	14
	GND
	WRITEn
	BR2
	AM5
	+3.3V

	15
	RsvBus
	GND
	BR3n
	A23
	GA3n

	16
	GND
	DTACKn
	AM0
	A22
	+3.3V

	17
	RsvBus
	GND
	AM1
	A21
	GA4n

	18
	GND
	Asn
	AM2
	A20
	+3.3V

	19
	RsvBus
	GND
	AM3
	A19
	RsvBus

	20
	GND
	IACKn
	GND
	A18
	+3.3V

	21
	RsvBus
	IACKINn
	SERA
	A17
	RsvBus

	22
	GND
	IACKOUTn
	SERB
	A16
	+3.3V

	23
	RsvBus
	AM4
	GND
	A15
	RsvBus

	24
	GND
	A01
	IRQ7n
	A14
	+3.3V

	25
	RsvBus
	A02
	IRQ6n
	A13
	RsvBus

	26
	GND
	A03
	IRQ5n
	A12
	+3.3V

	27
	RsvBus
	A04
	IRQ4n
	A11
	LI/In

	28
	GND
	A05
	IRQ3n
	A10
	+3.3V

	29
	RsvBus
	A06
	IRQ2n
	A09
	LI/On

	30
	GND
	A07
	IRQ1n
	A08
	+3.3V

	31
	RsvBus
	-12 VDC
	+5 VSTDBY
	+12 VDC
	GND

	32
	GND
	+5 VDC
	+5 VDC
	+5 VDC
	VPC


2.2.1.6 VMEBUS J2/P2 CONNECTOR PIN ASSIGNMENTS

A 96 pin connector will be used for P2/J2 because there are no proposed connections to Row D and Row Z.

Table 3 - VME P2 Signals

	VME I/O: P2/J2 Pin Assignments (VME64x)

	Pin
	Row Z
	Row A
	Row B
	Row C
	Row D

	1
	UsrDef
	UsrDef
	+5 VDC
	UsrDef
	UsrDef

	2
	GND
	UsrDef
	GND
	UsrDef
	UsrDef

	3
	UsrDef
	UsrDef
	RETRYn
	UsrDef
	UsrDef

	4
	GND
	UsrDef
	A24
	UsrDef
	UsrDef

	5
	UsrDef
	UsrDef
	A25
	UsrDef
	UsrDef

	6
	GND
	UsrDef
	A26
	UsrDef
	UsrDef

	7
	UsrDef
	UsrDef
	A27
	UsrDef
	UsrDef

	8
	GND
	UsrDef
	A28
	UsrDef
	UsrDef

	9
	UsrDef
	UsrDef
	A29
	UsrDef
	UsrDef

	10
	GND
	UsrDef
	A30
	UsrDef
	UsrDef

	11
	UsrDef
	UsrDef
	A31
	UsrDef
	UsrDef

	12
	GND
	UsrDef
	GND
	UsrDef
	UsrDef

	13
	UsrDef
	UsrDef
	+5 VDC
	UsrDef
	UsrDef

	14
	GND
	UsrDef
	D16
	UsrDef
	UsrDef

	15
	UsrDef
	UsrDef
	D17
	UsrDef
	UsrDef

	16
	GND
	UsrDef
	D18
	UsrDef
	UsrDef

	17
	UsrDef
	UsrDef
	D19
	UsrDef
	UsrDef

	18
	GND
	UsrDef
	D20
	UsrDef
	UsrDef

	19
	UsrDef
	UsrDef
	D21
	UsrDef
	UsrDef

	20
	GND
	UsrDef
	D22
	UsrDef
	UsrDef

	21
	UsrDef
	UsrDef
	D23
	UsrDef
	UsrDef

	22
	GND
	UsrDef
	GND
	UsrDef
	UsrDef

	23
	UsrDef
	UsrDef
	D24
	UsrDef
	UsrDef

	24
	GND
	UsrDef
	D25
	UsrDef
	UsrDef

	25
	UsrDef
	UsrDef
	D26
	UsrDef
	UsrDef

	26
	GND
	UsrDef
	D27
	UsrDef
	UsrDef

	27
	UsrDef
	UsrDef
	D28
	UsrDef
	UsrDef

	28
	GND
	UsrDef
	D29
	UsrDef
	UsrDef

	29
	UsrDef
	UsrDef
	D30
	UsrDef
	UsrDef

	30
	GND
	UsrDef
	D31
	UsrDef
	UsrDef

	31
	UsrDef
	UsrDef
	GND
	UsrDef
	GND

	32
	GND
	UsrDef
	+5 VDC
	UsrDef
	VPC


FRONT END PREAMP AND DIGITIZER

Each GRETINA Digitizer electronics board has 10 front-end input channels.  The data signals from the detector amplifiers are transmitted to the digitizer front-end inputs on differential wire pairs in a mass terminated cable assembly.  The receptacle connector on the Digitizer is a 100-pin double density D type with 75-mil contact spacing, manufactured by ITT/Cannon, part number 2DD100SBRP
.  The received input signal is buffered with a true differential input amplifier before the signal is digitized using a 14-bit 100Msps Analog to Digital Converter.  The digitized data of all channels is read out simultaneously by a controller/processor FPGA at 200Mbytes/sec/channel.

2.2.1.7 ANALOG INPUT

The analog inputs to the digitizer front-end amplifiers are true differential, high impedance type.   A block diagram is shown in Figure 3. 


[image: image3.wmf]GRETINA DIGITIZER CHANNEL FRONT END

DIGITIZER

100Msps ADC 

A IN 

-

A IN +

AD812

AD812

AD8138

Z

Comp


Figure 3: DIGITIZER FRONT END
The pin assignments for the analog input signals are defined in Table 4.  The labeling for the signals is as follows:

CH1P: Channel 1, positive polarity input signal

CH1N: Channel 1, negative polarity input signal

SHIELD1: Internal cable shield for Channel 1 wire pair

CH2P: Channel 2, positive polarity input signal

CH2N: Channel 2, negative polarity input signal

SHIELD2: Internal cable shield for Channel 2 wire pair

………

CH9P: Channel 9, positive polarity input signal

CH9N: Channel 9, negative polarity input signal

SHIELD9: Internal cable shield for Channel 9 wire pair

CH10P: Channel 10, positive polarity input signal

CH10N: Channel 10, negative polarity input signal

SHIELD10: Internal cable shield for Channel 10 wire pair

SP1: Auxiliary wire pair connections and shield

SP2: Auxiliary wire pair connections and shield

GND: Board Ground

Table 4 - Digitizer Input Pin Assignments

	Digitizer Input Interface 

(100pin ITT Cannon Double Density D Connector, 2DD100SBRP)

	Pin
	Description
	Pin
	Description
	Pin
	Description
	Pin
	Description

	1
	SHIELD10
	27
	SHIELD10
	
	
	
	

	2
	CH10P
	28
	CH10N
	52
	SP2
	76
	SP2

	3
	GND
	29
	GND
	53
	GND
	77
	SP2

	4
	GND
	30
	GND
	54
	SHIELD9
	78
	SHIELD9

	5
	GND
	31
	GND
	55
	CH9N
	79
	CH9P

	6
	SHIELD8
	32
	SHIELD8
	56
	GND
	80
	GND

	7
	CH8P
	33
	CH8N
	57
	GND
	81
	GND

	8
	GND
	34
	GND
	58
	GND
	82
	GND

	9
	GND
	35
	GND
	59
	SHIELD7
	83
	SHIELD7

	10
	GND
	36
	GND
	60
	CH7N
	84
	CH7P

	11
	SHIELD6
	37
	SHIELD6
	61
	GND
	85
	GND

	12
	CH6P
	38
	CH6N
	62
	GND
	86
	GND

	13
	GND
	39
	GND
	63
	GND
	87
	GND

	14
	GND
	40
	GND
	64
	SHIELD5
	88
	SHIELD5

	15
	GND
	41
	GND
	65
	CH5N
	89
	CH5P

	16
	SHIELD4
	42
	SHIELD4
	66
	GND
	90
	GND

	17
	CH4P
	43
	CH4N
	67
	GND
	91
	GND

	18
	GND
	44
	GND
	68
	GND
	92
	GND

	19
	GND
	45
	GND
	69
	SHIELD3
	93
	SHIELD3

	20
	GND
	46
	GND
	70
	CH3N
	94
	CH3P

	21
	SHIELD2
	47
	SHIELD2
	71
	GND
	95
	GND

	22
	CH2P
	48
	CH2N
	72
	GND
	96
	GND

	23
	GND
	49
	GND
	73
	GND
	97
	GND

	24
	GND
	50
	GND
	74
	SHIELD1
	98
	SHIELD1

	25
	SP1
	51
	SP1
	75
	CH1N
	99
	CH1P

	26
	SP1
	
	
	
	
	100
	GND


2.2.1.8 FRONT END PREAMP

High impedance input buffers are used to allow the cable termination resistors to dominate the impedance matching of the transmission line and to improve isolation while allowing signal balance on the wire pair.  The full-scale amplitude at the input of the front-end amplifier is +/- TBD.  

2.2.1.9 CENTRAL CONTACT REPEATER OUTPUT

The central contact of the detector shall be connected to the Channel 10 wire pair of the Master Digitizer that is connected to a differential repeater amplifier.  The repeater amplifier is used to buffer the Central Contact signal for connection to a front panel LEMO connector.  This allows the Central Contact signal to be viewed externally during run-time.

2.2.1.10 DIGITIZER CIRCUIT

The Digitizer circuit will utilize a 100Msps Analog Devices AD6645ASQ-105 ADC for each input channel.  The Main FPGA reads out each digitizer channel ADC simultaneously and writes the data to an internal circular buffer in real time.

2.2.2 EVENT FIFO

The Main FPGA writes full events into the Event FIFO after receiving a valid command from the trigger system.   The Event FIFO is implemented using two 256Kx18, 7ns FIFO memory chips.  Write control is managed solely by the Main FPGA while read control is managed by the Main FPGA through the VME interface.

2.2.3 MAIN FPGA  

Figure 1 shows the block diagram of the Digitizer.  The main FPGA is responsible for control the data acquisition.  It uses a Spartan 3 FPGA XC3S5000 (5M gates, 104 Block RAMs) from Xilinx.  The configuration memory has a minimum of 16 Mbits and is accessible from VME for in system reconfiguration.

Four Digitizers instrument one crystal.  One of these modules is configured as master and is connected to the crystal’s central contact.  The other three modules are configured as slaves. The VHDL code for all the 4 modules will be identical which means that we are going to have one part of the code that is the master and another part of the code that is the slave. The master code will be enabled only for the modules that are working as a master while the slave code will be enable in all the modules.  

Figure 4 shows a block diagram of the main FPGA.  All the FPGA logic is enabled when the module is configured as a master.  However, when the module is configured as a slave, the master logic is disabled.  The front bus logic receives commands and information from the front bus and either executes the command or sends back the information requested.  The VME control block does the communication between the main FPGA and the VME FPGA. This means that the main FPGA access the VME bus only through the VME FPGA.  The DAC control block is responsible for retrieving information from the channels, processing it, and controlling the DAC chips.  VME registers will configure the processing.  The channels block communicates with the FIFO interface.  The FIFO interface reads data out of the channel module, and during readout transfers data to the VME control.

[image: image4.wmf]ADC

ADC

ADC

ADC

ADC

ADC

ADC

ADC

ADC

ADC

FIFO

DAC

VME

FPGA

FB

Buffer

SD Chip

FIFO

Interface

DAC

Control

VME

Control

Front Bus

Logic

Master

Logic

Master

LogicEnable

C.C. LED

C.C. 

PileUp

EN

Channels


Figure 4: Main FPGA Block Diagram 
The channels block implements all the following DSP algorithms: 

· Leading Edge Discrimination: 

· yn = xn - xn-k (differentiation)

· yn = (xn + xn-2) + xn-1<<1  ((4, Gaussian filtering)

· Threshold comparison ( LED time

· Constant Fraction Discrimination: 

· yn = xn - xn-k (differentiation) 

· yn = (xn + xn-2) + xn-1<<1  ((2, Gaussian filtering)

· yn = xn-k- fxn  (constant fraction, f is an attenuation factor)

· Zero crossing comparison ( CFD time

· Trapezoidal filter and energy determination (V.T Jordanov, G.F. Knoll, NIM A345 (1994) 337-345)

· yn = yn-1 + ( (xn + xn-2m-k) )– (xn-m + xn-m-k) )
· Maximum tracking ( energy

· Pole-Zero correction
· yn = xn + In /t (where t is the pre-amplifier time constant)
· In = In-1 +  xn 

The master logic block will be enabled by the master logic enable signal.  This signal will be sent by the VME FPGA and will indicate to the main FPGA if it is a master or slave module, which will be determined by a VME register.

Table 6 – Master Logic Enable.

	Slot
	Master Enable
	Board type

	1
	1
	Master

	2
	0
	Slave

	3
	0
	Slave

	4
	0
	Slave


The master logic uses the front panel bus to control the four modules connected to a crystal.  Also, it interfaces with the serializer/deserializer (SerDes), which communicates with the TTCL.  Figure 5 shows the block diagram of the front bus master logic. 


[image: image5.wmf] 

SerDes

 

Chip

 

MasteSlave

 

Logic

 

Main Logic

 

TS Memory

 

TTCL

 

Fast Data

 

CH9

 

CC LED

 

CC 

 

PileUp

 

Master

 

Front Bus

 

Logic

 

SerDes

 

Logic

 


Figure 5: Block Diagram of the Front Bus Master Logic
The main logic block coordinates all the others block and interprets the TTCL commands.  The SerDes logic is responsible for reading and validating the data package received from the TTCL.  This data package is described in [3].  Once it validates the data package, it sends it to the main logic that will then execute the commands required.  It is also responsible for sending the information back to the TTCL as well as the fast data.  This other data package structure is described in [1].  In order to be able to send the fast data the central contact leading edge (LED) discriminator and the central contact pile up signals are connected directly to this block. The time stamp memory is a set of registers that store all the time stamps that were sent to the TTCL and still point to valid raw data inside the circular buffer.

2.2.3.1 MODES OF OPERATION

The main FPGA will operate in four different modes (see Table 7).  When a gamma ray interacts with the crystal, some of the ADC inputs will detect a leading edge (through the LED algorithm).  These digitizers will then set a special wired-or line in the front bus (refer to Section 2.2.6) to share the information that an event has taken place.
Table 7 - Master Logic Enable.

	Mode
	Board type

	00
	Internal

	01
	External

	10
	Internal with validation

	11
	TTCL


2.2.3.1.1 INTERNAL MODE

In the internal mode all the channels may start an event. But in order for a channel to be able to start an event it needs to be enabled. This will be done by one VME register.

Once a LED of any enabled channel happens then it will set the FB_LED line in the front bus that triggers all the channels. This will start the timers that will latch the results of the DSP algorithms (Energy, CFD, etc). Once all of these results are done, they will be stored in the Header Memory with index 0. After that they will be transferred to the pre-buffer memory. 

Finally this data will be readout by the FIFO interface. The FIFO interface is shared with all the 10 channels. It will first add the extra slave or master header (depending if the board is a slave or a master board) and then it will start reading out every channel data package. 

This mode can be used with one single board or one full crate without TTCL communication. In this mode it is not expected to have any validation so all the events will be readout.

2.2.3.1.2 EXTERNAL MODE

This mode is very similar to the internal mode. The difference between them is that in this mode the signal that set the FB_LED line in the front bus that triggers all the channels is not set by the LED of one of the channels but by one of the Auxiliary inputs.

2.2.3.1.3 INTERNAL WITH VALIDATION MODE

This mode is also very similar to the internal mode. The difference for this mode is that in it the results of the DSP algorithms and the raw data will not be automatically transferred to the pre-buffer memory. They will be waiting for a validation signal. The validation signal will be connected to one of the auxiliary inputs.

2.2.3.1.4 TTCL MODE

While operating in TTCL Mode the crystal crate will be connected through the serdes of the master board to the TTCL.

An event will start when a LED fires form the Central Contact (CC). The CC will be always connected to the Master Digitizer Board on channel 10. The CC LED information will be sent to all the channels in the four boards through a special line in the Front Bus and it will also be sent to the TTCL through the fast data. 

Once every channel received the CC LED information they will start the timers that will latch the results of the DSP algorithms (Energy, CFD, etc). Once all of these results are done, they will be stored in a Header Memory specifically defined to this event by the Master Digitizer Board. This means that all the channels are going to be storing the information of one event in the same Header Memory index.

The master board will be calculating the low energy of the CC and will request from the slave board and from itself the hit pattern. When both information are available in the Master Digitizer Board it will wait the TTCL to demand slow data from the digitizer boards, and then it will send these information to the TTCL. The Master Digitizer Board will also store the Time Stamp and the Header Memory index where the results of this event are stored. Once all the steps above are completed the boards are ready to handle another event or to receive a readout command.

Once the TTCL makes a decision that an event needs to be readout it will send to the digitizer boards a readout command and the time stamp of this event. The Master Digitizer Board will then compare this time stamp with all the time stamps that it has stored and are still valid. Valid means that the raw data are still available inside the circular buffer. If the Master Digitizer Board finds a match it will send a readout command to all channels informing which Header Memory needs to be readout.

Once the channels receive the readout command they will transfer the information from the Header Memory and the raw data to the pre-buffer memory. Then they will flag that this pre-buffer is ready to be sent to the external FIFO. The FIFO interface is shared with all the 10 channels. It will first add the extra slave or master header (depending if the board is a slave or a master board) and then it will start reading out every channel data package. This concludes the life cycle on an event that was readout.

If an event is not going to be readout because the TTCL does not sends a readout command for that particular event it will be flushed out the Header Memory after its corresponding data is over written in the circular buffer.

2.2.3.2 CHECKING SYNCHRONISM

Every channel has its own timer that generates the time stamp. So it is necessary to verify if all channels are in sync. Since the sync command will be sent every 2us and we don’t want to overload the front bus communication we are going to divide this task in two.

Every time the Master Digitizer board receives the sync command it will check the TTCL time stamp against the Central Contact time stamp, if both match then no error flag will be set. This is going to be done at a 2us second rate. 

In a slower rate the Master Digitizer Board will be checking the Central Contact time stamp against each channel. This is going to be done by sending a command that latch the current status of all channels as a broadcast command through the Front Bus. Then the Master Digitizer Board will broadcast the CC TS and each channel will compare the CC TS with its latched TS. If there is a mismatch the channel will raise an error flag.

The Master Digitizer Board will send a command to get the status of each board. If there is any error the Master Board will then signalize to the TTCL. No action besides send this information to the TTCL will be done. If a TS reset needs to be done, the TTCL will send specifically a command asking for that. Otherwise this flag will be set forever.

2.2.3.3 INITIALIZATION 

2.2.3.3.1 SYSTEM INITIALIZATION

All Digitizers in a crate will behave as slave boards after start up or reset.  The VME host can set a register in the Main FPGA that will determine if a board is a Master or Slave Digitizer.  After the VME host determines which board is a Master, the Master Logic Block can be enabled in one Digitizer per crate.  

2.2.3.3.2 INITIALIZATION OF TTCL COMMUNICATION

During initialization and reset the SERDES link will be down for some period of time while it gets lock. After locking the system will start receiving clock with a 20 ns period.

The TTCL will keep the SYNC signal active until its lock signal goes high and this means that the Master Digitizer Board will start receiving nine MSB equals to ‘0’ and nine LSB equals to ‘1’.  

Once the main FPGA has clock and the system reset goes low, then the Master Digitizer Board will be able to start sending data to the TTCL. Initially it will send sync data allowing the SERDES chip to get lock. After the TTCL SERDES lock then it will start sending real data down to the Master Digitizer Board.

The SD block of the Master Digitizer Board will get out the sync mode once it receives the first data that is different from the sync data. After that a different command will be sent to the main logic of the Master Logic (Figure 5) every 5 clocks.

If the lock signal goes down or a reset happens then the SD Block of the Master Digitizer Board will go back to the sync mode and it will be waiting for the same procedure to happen again in order to start receiving commands.

2.2.3.4 CLOCK MANAGEMENT 

The main FPGA has three clock inputs. One of them comes from the VME FPGA and is always present. The second one is the system clock source that comes either from the TTCL link or from the Front Bus. And the third one is an external clock from the auxiliary IOs.

When one board is working by itself the clock may come from the VME FPGA or from the auxiliary IOs. When the boards are working as an independent crate the slave boards will receive their clock through the Front Bus and the Main Digitizer Board will receive its clock from the VME FPGA or the auxiliary IOs.

When working on the TTCL mode the Main Digitizer Board will receive its clock from the recovered clock of the TTCL link and will spread it through the front bus to all the slave boards.

2.2.3.5 OUTPUT DATA FORMAT

The output data package will contain three types of header and the raw data that was read from the ADC. The structure of the data package can be seen on Figure 6
[image: image6.wmf]Slave Header

Package Header ch0

Raw Data

Package Header ch8

Raw Data

.

.

.

.

.

.

Slave Header

Package Header ch0

Raw Data

Package Header ch9

Raw Data

.

.

.

.

.

.

Master Header


Figure 6: (a) Data package from a slaver board. (b) Data package from a master board.

The Package header contains the information that belongs to the channels. Its structure is shown in Figure 10.

	0
	1
	2
	3
	4
	5
	6
	7
	8
	9
	10
	11
	12
	13
	14
	15
	16
	17
	18
	19
	20
	21
	22
	23
	24
	25
	26
	27
	28
	29
	30
	31

	Board ID1
	Packet length (header included)

	LED / external timestamp bit 0 - 15*
	LED / external timestamp bit 16 – 31*

	LED / external timestamp bits 32 - 48*
	Energy bits 0 – 15

	Energy bit 16 – 24
	x
	x
	x
	S
	E
	C
	P
	CFD Timestamp bit 0 – 15

	CFD Timestamp bit 16 - 31
	CFD Timestamp bit 32 – 47

	CFD point 1 bit 0 - 15
	CFD point 1 bit 16 – 31

	CFD point 2 bit 0 - 15
	CFD point 2 bit 16 – 31

	Raw data point 0 (Sign extended)
	Raw data point 1 (Sign extended)

	Raw data point 2 (Sign extended)
	Raw data point 3 (Sign extended)

	.

.

.
	.

.

.


Figure 7: Data Package.

*The TS is updated at 20ns.

Notes:

S is for the sign of the LED crossing. (1 is negative).

E is for external trigger flag (the timestamp is external and LED and CFD are not valid).

C is for CFD valid indicating that a CFD crossing occurred.

P is for pileup flag indicating that the energy is corrupted.

1Board ID will be composed of 4 bits for channel ID and 12 bits for the board serial number.

The slave header contains information from the event that is common to all the channels on that specific board.  All Digitizers in a crate produce the Slave Header.  The Master Digitizer will add a top-level header to the data packet that is used to identify the start of a new event from a crystal.  
2.2.3.6 MAIN FPGA REGISTERS

The address shown in the tables that describe the registers are only the bits 9 down to 2. This makes the addresses here compatible with the address written inside the VHDL code.

	Description
	Address
	Access
	Width

	Board ID
	0x00
	R
	32

	Value is the board serial number that is given by the switch configuration. 
	Bit Value

	
	1
	0

	  Bits[11: 0]: Board Id
	Value

	  Bits[31:12]: TBD 
	


	Description
	Address
	Access
	Width

	Programming Done
	0x01
	R
	32

	Value is the board serial number that is given by the switch configuration. 
	Bit Value

	
	1
	0

	  Bits[9: 0]: Programming status of the channels
	
	

	  Bit 10: Programming status of the debug module 
	
	

	  Bit 11: Programming status of the DAC module 
	
	

	  Bit 12: and FIFO 0 EF flag
	
	

	  Bit 13: and FIFO 1 EF flag
	
	

	  Bit 14: and FIFO 0 PAE flag
	
	

	  Bit 15: and FIFO 0 HF flag
	
	

	  Bit 16: and FIFO 0 PAF flag
	
	

	  Bit 17: and FIFO 0 FF flag and FIFO 1 FF flag
	
	

	  Bits[31:18]: TBD
	


	Description
	Address
	Access
	Width

	External Window
	0x02
	W
	32

	External validation window length. Value at reset is 0x07FF.
	Bit Value

	
	1
	0

	  Bits[10: 0]: External window length
	Value

	  Bits[31:11]: TBD 
	


	Description
	Address
	Access
	Width

	Pileup Window
	0x03
	W
	32

	Pileup window length. Value at reset 0x0400 (10us).
	Bit Value

	
	1
	0

	  Bits[10: 0]: Pileup window length
	Value

	  Bits[31:11]: TBD 
	


	Description
	Address
	Access
	Width

	Noise Window
	0x04
	W
	32

	Noise window length. Value at reset 0x0040 (640ns).
	Bit Value

	
	1
	0

	  Bits[ 6: 0]: Noise window length
	Value

	  Bits[31:11]: TBD 
	


	Description
	Address
	Access
	Width

	External trigger sliding length
	0x05
	W
	32

	Length before we read the energy when we operate in external trigger. Value at reset 0x01C2 (4.5us)
	Bit Value

	
	1
	0

	  Bits[10: 0]: External trigger sliding length
	Value

	  Bits[31:11]: TBD 
	


	Description
	Address
	Access
	Width

	Collection time
	0x06
	W
	32

	Collection time maximum length (length of the flat top in the trapezoid).

Value at reset 0x01C2 (4.5us)
	Bit Value

	
	1
	0

	  Bits[8: 0]: Collection time
	Value

	  Bits[31:9]: TBD 
	


	Description
	Address
	Access
	Width

	Integration time
	0x07
	W
	32

	Integration time length (length of one side of the trapezoid). 

Value at reset 0x01C2 (4.5us)
	Bit Value

	
	1
	0

	  Bits[8: 0]: Integration time
	Value

	  Bits[31:9]: TBD 
	


	Description
	Address
	Access
	Width

	Control/Status channel #Number
	*
	RW
	32

	This register address is the only one that is read/write.
	Bit Value

	
	1
	0

	  Bit 0: START/STOP. This bit must be set to one for the channel to operate. Value at reset is 0. Read/Write.
	
	

	  Bit 1: Debug Mode. This bit when set to 1 puts the channel in debug mode operation using internal data. Value at reset 0. (Read/Write)
	
	

	  Bit 2: Pile-up drop-out enable. When this bit is set to one, the channel drops any event that occurs in a pileup window else it just notifies pileup through a flag in

the event header. Value at reset 1. Read/Write.
	
	

	  Bits [4:3]: Trigger mode. We have the following configuration:

00: Internal mode

01: External mode

10: Internal with Validation mode

11: TTCS Decision mode

Value a reset: 00. Read/Write.
	
	

	  Bit 5: CFD Tap delay setup. If set to 0, the tap delay is filling-up and no event can be acquired. Read-only.
	
	

	  Bit 6: Tap delay 1 setup. If set to 0, the tap delay is filling-up and no event can be acquired. Read-only.
	
	

	  Bit 7: Tap delay 2 setup. If set to 0, the tap delay is filling-up and no event can be acquired. Read-only.
	
	

	  Bit 8: Tap delay 3 setup. If set to 0, the tap delay is filling-up and no event can be acquired. Read-only.
	
	

	  Bit 9: Tap delay 4 setup. If set to 0, the tap delay is filling-up and no event can be acquired. Read-only.
	
	

	  Bits [11:10]: Polarity validation. Value at reset 11.

01: only positive trigger considered

10: only negative trigger considered

11: both triggers considered

00: no trigger considered (the LED still fires

externally though, as opposed to

START/STOP) Read/Write.
	
	

	  Bits[14:12]: TBD
	
	

	  Bit 15: Pre-buffer Ready flag
	
	

	  Bits[31:16]: TBD 
	


· There is one of this registers for each channels and their address are specified by the next table

	Description
	Address
	Access
	Width

	Control/Status channel 0
	0x10
	RW
	32

	Control/Status channel 1
	0x11
	RW
	32

	Control/Status channel 2
	0x12
	RW
	32

	Control/Status channel 3
	0x13
	RW
	32

	Control/Status channel 4
	0x14
	RW
	32

	Control/Status channel 5
	0x15
	RW
	32

	Control/Status channel 6
	0x16
	RW
	32

	Control/Status channel 7
	0x17
	RW
	32

	Control/Status channel 8
	0x18
	RW
	32

	Control/Status channel 9
	0x19
	RW
	32


	Description
	Address
	Access
	Width

	LED Threshold
	*
	W
	32

	This is the LED threshold. Only bits 17 to 0 are used. This is an unsigned value and the three lower bits are extra precision bit. (The dot is between bit 3 and 2).

The value at reset is 0x7FFF (full range so that no event should trigger). It is internally converted to a signed value depending on how the sign of the current sample is.

Write-only.
	Bit Value

	
	1
	0

	  Bits[17: 0]: LED Threshold
	Value

	  Bits[31:18]: TBD 
	


* There is one of this registers for each channels and their address are specified by the next table.

	Description
	Address
	Access
	Width

	Control/Status channel 0
	0x20
	W
	32

	Control/Status channel 1
	0x21
	W
	32

	Control/Status channel 2
	0x22
	W
	32

	Control/Status channel 3
	0x23
	W
	32

	Control/Status channel 4
	0x24
	W
	32

	Control/Status channel 5
	0x25
	W
	32

	Control/Status channel 6
	0x26
	W
	32

	Control/Status channel 7
	0x27
	W
	32

	Control/Status channel 8
	0x28
	W
	32

	Control/Status channel 9
	0x29
	W
	32


	Description
	Address
	Access
	Width

	CFD Parameters
	*
	W
	32

	Bit 12-7: CFD delay. Value at reset is 0x3F giving a delay of 630ns. Write-only

Bit 6-5: CFD fraction. It is a value indicating what fraction is used (see table 9 for values). Value at reset is “00” giving a fraction of 0.5. Write-only

Bit 4-0: CFD threshold. It is an unsigned value of the CFD threshold. The dot is after bit 0. Value at reset is 0x10 (160kev). Write-only
	Bit Value

	
	1
	0

	  Bits[ 4: 0]: CFD threshold
	Value

	  Bits[ 6: 5]: CFD fraction
	Value

	  Bits[12: 7]: CFD delay
	Value

	  Bits[31:13]: TBD 
	


* There is one of this registers for each channels and their address are specified by the next table.

	Description
	Address
	Access
	Width

	Control/Status channel 0
	0x30
	W
	32

	Control/Status channel 1
	0x31
	W
	32

	Control/Status channel 2
	0x32
	W
	32

	Control/Status channel 3
	0x33
	W
	32

	Control/Status channel 4
	0x34
	W
	32

	Control/Status channel 5
	0x35
	W
	32

	Control/Status channel 6
	0x36
	W
	32

	Control/Status channel 7
	0x37
	W
	32

	Control/Status channel 8
	0x38
	W
	32

	Control/Status channel 9
	0x39
	W
	32


	Description
	Address
	Access
	Width

	Raw data sliding length
	*
	W
	32

	Delay before we start retrieving raw data. Only bits 10 to 0 are used. Value at reset is 0x01C2 (4.5us). Write-only
	Bit Value

	
	1
	0

	  Bits[10: 0]: Raw data sliding length
	Value

	  Bits[31:11]: TBD 
	


* There is one of this registers for each channels and their address are specified by the next table.

	Description
	Address
	Access
	Width

	Control/Status channel 0
	0x40
	W
	32

	Control/Status channel 1
	0x41
	W
	32

	Control/Status channel 2
	0x42
	W
	32

	Control/Status channel 3
	0x43
	W
	32

	Control/Status channel 4
	0x44
	W
	32

	Control/Status channel 5
	0x45
	W
	32

	Control/Status channel 6
	0x46
	W
	32

	Control/Status channel 7
	0x47
	W
	32

	Control/Status channel 8
	0x48
	W
	32

	Control/Status channel 9
	0x49
	W
	32


	Description
	Address
	Access
	Width

	Raw data window
	*
	W
	32

	Length of raw data retrieved. Only bits 9 to 0 are used. Value at reset is 0x32 (500ns). Write-only
	Bit Value

	
	1
	0

	  Bits[ 9: 0]: Raw data length
	Value

	  Bits[31:10]: TBD 
	


* There is one of this registers for each channels and their address are specified by the next table.

	Description
	Address
	Access
	Width

	Control/Status channel 0
	0x50
	W
	32

	Control/Status channel 1
	0x51
	W
	32

	Control/Status channel 2
	0x52
	W
	32

	Control/Status channel 3
	0x53
	W
	32

	Control/Status channel 4
	0x54
	W
	32

	Control/Status channel 5
	0x55
	W
	32

	Control/Status channel 6
	0x56
	W
	32

	Control/Status channel 7
	0x57
	W
	32

	Control/Status channel 8
	0x58
	W
	32

	Control/Status channel 9
	0x59
	W
	32


	Description
	Address
	Access
	Width

	Debug data buffer address
	0x60
	W
	32

	Used by debug module (see section XIII of the Gretina VHDL modules description document).
	Bit Value

	
	1
	0

	  Bits[31: 0]: 
	Value


	Description
	Address
	Access
	Width

	Debug data buffer data
	0x61
	W
	32

	Used by debug module (see section XIII of the Gretina VHDL modules description document).
	Bit Value

	
	1
	0

	  Bits[31: 0]: 
	Value


	Description
	Address
	Access
	Width

	LED Flag window
	0x62
	W
	32

	Number of clock cycles the LED flag is kept raise on the Fast Data. Range varies from 1 (10ns) to 1024 (~1us).
	Bit Value

	
	1
	0

	  Bits[31: 0]: 
	Value


	Description
	Address
	Access
	Width

	Internal Mode operation Mask
	0x63
	W
	32

	When operating in internal mode this register enable the channels that can write on the fast signal on the FB Leading edge detected.
	Bit Value

	
	1
	0

	  Bits[10: 0]: Channels that are enabled
	Value

	  Bits[31:10]: TBD
	


	Description
	Address
	Access
	Width

	Front Bus Enabled
	0x64
	W
	32

	This is only used if the board in not working on the TTCL mode. If the front bus is enable means that we have a full crate (four boards, one master and three slaves) otherwise the board is working by itself.
	Bit Value

	
	1
	0

	  Bit 0: Front Bus enabled 
	Value

	  Bits[31: 1]: TBD
	Value


2.2.4 DAC OUTPUTS

The GRETINA Digitizer will utilize 2 general-purpose 100MHz Digital to Analog converters (DAC) that will have outputs available at the front panel of the module.  The design will be implemented with current output DACs to allow signal summing using a NIM module.  The board will be designed to allow the option of loading the circuit as a differential voltage output DAC if required.  The option must be selected at the time that the board is assembled and will not be user selectable in the system.  Signals required by the DACs: Clock and 8 bit data word.  (Analog Devices AD9748ACP)

2.2.5 FRONTBUS INTERFACE 

The Front Bus Interface will be used by the Master Digitizer to distribute the system clock and to communicate with the three Slave Digitizers in each VME crate.  The Master will set some control registers asking the boards to execute some task and will read and write from other registers in order to transfer information. The Master Digitizer is always connected to the Central Contact of the Crystal, so the time require to fetch hit patterns will not affect the ability to supply information to the “Fast Decision Trigger” algorithms.

The 50MHz System Clock is recovered from the Serial Data transmitted by the Global Trigger and Timing crate and then routed as a differential pair to all of the Digitizers over the Front Bus.  The 50MHz clock is multiplied to 100MHz in the Main FPGA and then used to digitize the input data for each channel.  

A block diagram of the system clock distribution circuit is shown in Figure 8.


[image: image7.wmf]GRETINA FRONTBUS CLOCK: BLOCK DIAGRAM

MAIN 

FPGA

FRONTBUS  IO

MDR20 

Connector

SER/DES   IO

BLOCK

50MHz RCLK

SYSTEM CLOCK SOURCE SELECT

50MHz SYSTEM CLK


Figure 8: FRONTBUS CLOCK DISTRIBUTION

Data Bus: 10 bit bi-directional parallel bus

a. Valid data latched on each clock transition

b. The data and command bus are implemented using standard LVT245 transceivers (Bus Hold and built in series termination resistors)

Address Bus: 8 bits wide  

Table 9 – Front Bus Base Address, Bits[7:6]

	Front Bus Base Address (Bits [7:6]
	Target Device for Read Access only

	00
	Digitizer 0 (Master)

	01
	Digitizer 1

	10
	Digitizer 2

	11
	Digitizer 3


For Read access, the Master activates one Digitizer per complete cycle.

Bit 5 is used to indicate a Broadcast function is active on the Front Bus.

Table 10 – Front Bus commands, Address Bits[5:0]

	Front Bus Command
	Command Name
	Command Description

	100000
	Write trigger data to slave.
	This command performs a read out operation and will send the Header Memory index in the data.

	1000001
	Reset TS
	All time stamps will be reset.

	1000010
	Send TS
	Send the time stamp that will be compared with the channel TS to verify if they are in sync

	1000011
	Latch status
	Latch the status of all channels in the same time, so it is possible to compare the TS of the channels.

	1000100
	Calibration Inject 
	

	
	
	


* This is an option if we want to have more then 10 bits of status for every board.

Leading Edge Detect Wired-OR signal – 1 bit

This signal is used to indicate that a detector crystal has detected a leading edge.

Table 11 - Front Bus Interface
	Front Bus Interface (40-pin MT Connector AMP)

	Pin
	NAME
	Description

	1
	CLK_IO_P
	FB Clock positive polarity

	2
	CLK_IO_N
	FB Clock negative polarity

	3
	GND
	Board Ground

	4
	GND
	Board Ground

	5
	DATA(0)
	FB Data Bus Bit 0

	6
	DATA(1)
	FB Data Bus Bit 1

	7
	DATA(2)
	FB Data Bus Bit 2

	8
	GND
	Board Ground

	9
	DATA(3)
	FB Data Bus Bit 3

	10
	DATA(4)
	FB Data Bus Bit 4

	11
	DATA(5)
	FB Data Bus Bit 5

	12
	GND
	Board Ground

	13
	DATA(6)
	FB Data Bus Bit 6

	14
	DATA(7)
	FB Data Bus Bit 7

	15
	DATA(8)
	FB Data Bus Bit 8

	16
	GND
	Board Ground

	17
	DATA(9)
	FB Data Bus Bit 9

	18
	ADDR(0)
	FB Address Bit 0

	19
	ADDR(1)
	FB Address Bit 1

	20
	GND
	Board Ground

	21
	ADDR(2)
	FB Address Bit 2

	22
	ADDR(3)
	FB Address Bit 3

	23
	ADDR(4)
	FB Address Bit 4

	24
	GND
	Board Ground

	25
	ADDR(5)
	FB Address Bit 5

	26
	ADDR(6)
	FB Address Bit 6

	27
	ADDR(7)
	FB Address Bit 7

	28
	GND
	Board Ground

	29
	READ
	FB Read Enable

	30
	WRITE
	FB Write Enable

	31
	STROBE
	FB Data Strobe

	32
	GND
	Board Ground

	33
	LE DETECT
	FB Leading Edge Detect (Wired OR)

	34
	SPARE
	Spare Pin

	35
	SPARE
	Spare Pin

	36
	SPARE
	Spare Pin

	37
	GND
	Board Ground

	38
	SPARE W-OR
	Spare Pin (Wired O R)

	39
	SP DIFF PAIR
	Spare Pin DIFF

	40
	SP DIFF PAIR
	Spare Pin DIFF


There will be two types of commands for the Front Bus, the read and write. The signal will be asserted on the falling edge of the clock and will be valid on the raising edge of the clock. The timing diagrams of these two commands are shown in Figure 13 and Figure 15. 

During a reset all the modules will start working as slave, so none of them are going to be driving the control signals. The Strobe*
 line is the signal that tells the board that a new command needs to be executed. In order to avoid any false commands from being executed right after a reset this line will have a pull-up resistor.  

[image: image8.wmf]DATA

CLK

Address

Strobe*

ZZ

Data

XX

ZZ

Address

XX

RNW

XX

XX


Figure 9: Front Bus communication timing diagram – master write operation.
[image: image9.wmf]DATA

CLK

Address

Strobe*

ZZ

Data

XX

ZZ

Address

XX

RNW

XX

XX


Figure 10: Front Bus communication timing diagram – master read operation.
2.2.5.1 FRONT BUS REGISTERS

	Description
	Address
	Access
	Width

	Slave Board Hit Pattern Register
	BASE ADDRESS + 0x00
	R
	10

	Hit pattern of the latest event that happened.
	Bit Value

	
	1
	0

	  Bits[9: 0]: Hit Pattern
	Value


	Description
	Address
	Access
	Width

	Slave Board Status Register
	BASE ADDRESS + 0x01
	R
	10

	Status of Slave Board, Read the status of every board. *Data contains the information about which status is going to be readout.
	Bit Value

	
	1
	0

	  Bits[9: 0]: Status (Bit definitions TBD)
	Value


	Description
	Address
	Access
	Width

	Time stamp 0
	BASE ADDRESS + 0x02
	R
	10

	These are part of the 48 bits of the time stamp that will be used the check the synchronism of the system.
	Bit Value

	
	1
	0

	  Bits[9: 0]: Time stamp bits [ 9: 0]
	Value


	Description
	Address
	Access
	Width

	Time stamp 1
	BASE ADDRESS + 0x03
	R
	10

	These are part of the 48 bits of the time stamp that will be used the check the synchronism of the system.
	Bit Value

	
	1
	0

	  Bits[9: 0]: Time stamp bits [19:10]
	Value


	Description
	Address
	Access
	Width

	Time stamp 2
	BASE ADDRESS + 0x04
	R
	10

	These are part of the 48 bits of the time stamp that will be used the check the synchronism of the system.
	Bit Value

	
	1
	0

	  Bits[9: 0]: Time stamp bits [29:20]
	Value


	Description
	Address
	Access
	Width

	Time stamp 3
	BASE ADDRESS + 0x05
	R
	10

	These are part of the 48 bits of the time stamp that will be used the check the synchronism of the system.
	Bit Value

	
	1
	0

	  Bits[9: 0]: Time stamp bits [39:30]
	Value


	Description
	Address
	Access
	Width

	Time stamp 4
	BASE ADDRESS + 0x06
	R
	10

	These are part of the 48 bits of the time stamp that will be used the check the synchronism of the system.
	Bit Value

	
	1
	0

	  Bits[7: 0]: Time stamp bits [47:40]
	Value

	  Bits[9: 8]: TBD
	Value


2.2.6 SERIALIZER/DESERIALIZER

The Digitizer receives Trigger commands, transmits event information and receives the Gretina System Clock through the Serializer/Deserializer circuit.  The 50 MHz system clock is generated by the Global Trigger and recovered by the SER/DES on each Master Digitizer.  The circuit design includes the National Semiconductor DS92LV18TVV, LVDS Serializer/Deserializer and a pre-emphasis IC used to boost the fidelity of the transmitted signals.  A custom, shielded RJ45 to 2mm hard Metric cable will be used to connect the Digitizer SER/DES with the Global trigger SER/DES.

More detailed documentation with regard to the SER/DES circuits can be found in these [1, 2, 3] documents:

2.2.7 AUXILLIARY INPUT INTERFACE

The Auxiliary Input Interface is used to link the Gretina Digitizers to other dectectors and experiments.  There are 16 inputs on the front panel connectors, and one is a dedicated external clock input that can be used to sync the Digitizer ADC circuits to another detector.  The connectors are standard headers with 0.1x0.1 spacing, and there is 1 ground per input pin.

2.2.8 100MHz CLOCK DISTRIBUTION

The MAIN FPGA will generate a 100MHz clock from the recovered 50MHz system clock and this clock will by distributed as an LVDS signal to all of the Analog Front Ends.  List Part -- State skews and delays
2.2.9 FRONT PANEL LED INDICATORS

The Digitizer Front Panel will include 18 dual color LED indicators.  The indicators are assigned as follows:

The MAIN FGPA controls the state of 13 front panel indicators.  There is 1 bi-color, green/yellow or green/red (10) per Digitizer Channel.


States:


Channel Off ( LED Off


Channel On ( GREEN Blinking 


Taking Data ( GREEN Solid


FIFO Readout ( YELLOW Solid


Rate OVR ( YELLOW Blinking

The other 3 indicators that are controlled by Main FPGA are user definable, and can be modified with changes to the MAIN FPGA firmware.

The VME FPGA controls the state of 2 LEDs, and 3 are connected directly to key power supply planes.


VME Activity ( YELLOW


FPGA READY ( GREEN


+3.3V DC ( GREEN


+5.0V DC ( GREEN


-5.0V DC ( GREEN

2.2.10 VOLTAGE MONITORS

The Digitizer has a voltage monitor chip on-board to monitor key power supply voltages in real time.  The monitor circuit includes an “All-OK” signal that is connected to the on-board power-on reset circuit and to a VME FPGA Status Registe There is also an under-voltage signal and an over-voltage signal that are also connected to the same VME FPGA register.  The monitored voltages are +3.3V, +5.0V, -5.0V, and +1.2V (Main FPGA).  The circuit is powered by the +12V supply, and if the 1.2V VME core voltage is faulty than the board will not be readable on the VME bus.

2.2.11 TEMPERATURE MONITORS

The Digitizer is design wit three temperature monitoring switches that trip if board temp exceed 70C.  These signal are all connected to a VME FPGA status register that can be monitored by the VME host.

3 REFERENCES
[1]Anderson, J. Link Specification: Trigger Input Data Link. v.1.3. Gretina note. August, 2006.

[2] Anderson, J. Link Specification: Trigger Module (Router & Master Trigger). v1.1. Gretina note. August, 2006.

[3] Anderson, J.; Zimmermann, S. Link Specification: Trigger Timing & Control Link. v1.5. Gretina note. August, 2006.

Gamma Ray Energy Tracking


In-Beam Nuclear Array















































� * means that the signal has its logic inverted.





_1216498587.ppt


GRETINA DIGITIZER BLOCK DIAGRAM

CH0 FE/ADC

DAC0 OUT

LED OUT

AUX IO IN

FRONTBUS  IO

SER/DES   IO

CC OUT

50MHz REF CLK

50MHz RCLK

CH1 FE/ADC

CH9 FE/ADC

DAC1 OUT





EVENT FIFO

















VME CTRL

& 

I/O





MAIN FPGA

XC3S5000-5FG900

(633 USER I/O)



100MHz CLOCK DIST



MUX








_1216825582.doc


SerDes







Chip







MasteSlave







Logic







Main Logic







TS Memory







TTCL







Fast Data







CH9







CC LED







CC 







PileUp







Master







Front Bus







Logic







SerDes







Logic












_1217020371.ppt


GRETINA DIGITIZER CHANNEL FRONT END

A IN -

A IN +

AD812

AD812

AD8138



DIGITIZER



100Msps ADC 







Z

Comp 








_1216501775.ppt


GRETINA FRONTBUS CLOCK: BLOCK DIAGRAM

FRONTBUS  IO

MDR20 Connector

SER/DES   IO

BLOCK

50MHz RCLK

SYSTEM CLOCK SOURCE SELECT

50MHz SYSTEM CLK



MAIN FPGA
















_1216498470.ppt


GRETINA VME INTERFACE BLOCK DIAGRAM

MAIN FPGA CONFIG. MEMORY (128/64/32 Mbits)

16

DATA

CTRL

BUSY

MAIN FPGA CONFIG

4

5

32

DATA

32

ADDR

6

AM

VME CTRL

7

50MHz REFERENCE CLOCK

DATA

ADDR

32

32

GADDR

6

VME SPEC

8

MEMORY MAP CTRL

3

SYSTEM MONITORING

6

IRQ

10

















VME BUS 

I/O





VME FPGA

XC3S400-4FG320

(221 USER I/O)



50MHz CLOCK DIST








