## Slide 1: Power Supply System
John Joseph
LBNL
Spectrometer Solenoid
Test Plan Review
Feb 17th 2012

## Slide 2: Simplified PS System Block Diagram
Power Supply Slow Control System
Quench Detect System
Slow Control Functions
  Ramp Rates
  Contactor switching for ramp cycles
  Status Monitoring
    Voltage
    Current
    Temperature
  Direct control of all PS Units
QDS Functions
  Real time Quench Detection
  Contactor switching for fast discharge
  Real time system status monitoring
Magnet Coil Power Supply
Rack
Power Supplies
for the
Center Coil &
Matching Coils

## Slide 3: Primary PS Rack Components and Loads

## Slide 4: Wang NMR System: DC SIGNAL PATH & Control

## Slide 5: LBNL/MICE: DC PATH [As Built]
Added 5 contactors to system
  HTS Lead protection during fast discharge

## Slide 6
Energy Absorber Assemblies
Plumbing for Water Cooling of Energy Absorbers
4 Energy Absorbers
  Required for slow ramp down of coils
  2 diode assemblies per absorber circuit
    1 DA always dissipating power into air & water
  2 water circuits available for cooling

## Slide 7: POWER SUPPLY RACK CONFIGURATION
MS2 300A PS
MS1 300A PS
EC1 60A PS
EC2 60A PS
CS 60A PS
EA - DA3
EA - DA4
EA - DA2
EA - DA1
K2
K1
K7
K8
K3
K6
FRONT VIEW

## Slide 8
Redesigned system to improve HTS Lead protection
  Additional contactors to isolate coils from power supplies during fast ramp down and enable in-cryo coil protection
  Adding 2nd Solid State Relay to handle increased Contactor pickup current
Reconfigured rack to optimize DC cabling
  Improved physical separation of AC and DC wires
  All Contactors physically close to associated supplies
Moved all AC wiring above water cooling circuits
Replacing 15 VDC PS with 24 VDC PS for interlock circuits
Replaced all control and monitoring wiring
  Bundled twisted pair
Summary of Changes Implemented @ LBL

## Slide 9
Requires 2nd Solid State Relay
  ~35A of in-rush when coils energized to close contactors
Replacing 15 VDC PS with 24 VDC PS
  Standard for Interlock systems
  Requested by RAL
Potential to stagger Contactor on signal to reduce instantaneous power peak

## Slide 10
Power Entry, Control, and Monitoring
TB1
208 VAC 3-PHASE
Power Entry
TB4
Status Monitoring
SSR, TB3, PS6
Interlock & Contactor Control
TB2
PS Voltage Monitoring
TB4
Status Monitoring
K1 - VDROP
K2 – VDROP
K3 – VDROP
K4 - VDROP
K5 - VDROP
K6 - VDROP
K7 - VDROP
K8 – VDROP
CONTACT STATUS
PS TEMP SENSORS

## Slide 11
Rewire Temp Sensors to move over-temp reaction control to Slow or Fast DAQ
  Overheating of Energy Absorbers should be slow ramp down not fast discharge
Add Emergency Off switch to front and back of racks
Contactor Aux Contacts indicate
  All Contactors closed
  One or more Contactors open
    Currently only indicates if all are open
Additional Changes per Collaboration Discussion

## Slide 12
Loads on 3 Phase 208 VAC
  300A Power Supply for the Center Coil (15A @ 3.3kW)
  300A Power Supply for Matching Coil 1 (80% efficiency min)
  300A Power Supply for Matching Coil 2 (80% efficiency min)
Loads on Single Phase 208 VAC
  6 Contactor Coils requiring 970VA for Pickup
  2 Contactor Coils requiring 245VA for Pickup
  24VDC, 60 Watt, Power Supply for system Interlock circuits
Loads on 120 VAC
  60A Trimming Power Supplies for the End Coils (Qty 2)
    850 VA per Unit
  Power Supply Controllers for the 300A Units (Qty 2)
    50 VA per Unit
Power Requirements for Test

## Slide 13
Complete remaining assembly tasks on the rack to be used at Wang
  Implement RAL recommendations
Complete the final Documentation package
  Final Schematics & Wire Lists
  Configuration Drawing
Low Current Functionality Testing of all System Components
  Avoid connecting water circuits
  Load PS outputs with shorting cables
Decision on Assembly of 2nd Power Supply Rack
  Requested to go directly to RAL
  New rack configuration is under consideration
Tasks to Complete