## Change Log
2011-11-29 | Internal Sensors CX1 and CX2 do not have calibration curves, they will be replaced by new sensors on the outside of the chamber. Output from these sensors will still be available on feedthrough H1.
Sensor SD18 and SD22 removed from the bore of the cold mass, not enough clearance and would have localy crushed the MLI
Sensor SD18 replaced on feedthrough V4 by new CX1 on outside of the chamber
Sensor SD22 replaced on feddthrough V4 by new CX2 on outside of the chamber
Sensor SD25 changed to LR type

## Sensor & Equipment Map
[16 empty] | Quality Checking Magnet 1 |  |  |  |  | Quality Checking Magnet 2
Connector | Type | Pin No | Pin Letter | Designation | Polarity / Position | Type | Part No | Cable Label | Serial Number | Placement | Comments | Monitor Type | Monitor | Input | Pin No | R V+V- | RI+I- | R V+I+ | R V-I- | Who | R V+V- | RI+I- | R V+I+ | R V-I- | Who
H1 | Ceramic | 1 | A | CX01 | +V | Cernox | Unknown | CX/1/H1/A/+V | Unknown | He - End Coil 1 | Feedthrough will be on the Fill / Vent tower | Lakeshore 218 | 5 | 5 | 4
| 10 Pin | 2 | B | CX01 | -V |  |  | CX/1/H1/B/-V |  |  | He can | Lakeshore 218 | 5 | 5 | 16
| 2.75" CF | 3 | C | CX01 | +I |  |  | CX/1/H1/C/+I |  |  |  | Lakeshore 218 | 5 | 5 | 3
|  | 4 | D | CX01 | -I |  |  | CX/1/H1/D/-I |  |  |  | Lakeshore 218 | 5 | 5 | 15
| (Old J1) | 5 | E | CX02 | +V | Cernox | Unknown | CX/2/H1/E/+V | Unknown | He - Match Coil1 |  | Lakeshore 218 | 5 | 6 | 7
|  | 6 | F | CX02 | -V |  |  | CX/2/H1/F/-V |  |  |  | Lakeshore 218 | 5 | 6 | 19
|  | 7 | G | CX02 | +I |  |  | CX/2/H1/G/+I |  |  |  | Lakeshore 218 | 5 | 6 | 6
|  | 8 | H | CX02 | -I |  |  | CX/2/H1/H/-I |  |  |  | Lakeshore 218 | 5 | 6 | 18
|  | 9 | J
|  | 10 | K
H2 | Ceramic | 1 | A | LHE01 | +V | Helium Level |  | LV/1/H2/A/+V | Unknown | He | Bottom Sensor | AMI 135
| 19 Pin | 2 | B | LHE01 | -V |  |  | LV/1/H2/B/-V |  |  |  | AMI 135
| 2.75" CF | 3 | C | LHE01 | -I |  |  | This cable is common to both sensor 1 and 2. Only 1 cable exiting the cold mass |  |  |  | AMI 135
|  | 4 | D | LHE01 | +I |  |  | LV/1/H2/D/+I |  |  | Feedthrough will be on the Fill / Vent tower He can | AMI 135
|  | 5 | E | LHE02 | +V | Helium Level |  | LV/2/H2/E/+V | Unknown | He | Top Sensor | AMI 135
| (Old J2) | 6 | F | LHE02 | -V |  |  | LV/2/H2/F/-V |  |  |  | AMI 135
|  | 7 | G | LHE02 | -I |  |  | This cable is common to both sensor 1 and 2. Only 1 cable exiting the cold mass |  |  |  | AMI 135
|  | 8 | H | LHE02 | +I |  |  | LV/2/H2/H/+I |  |  |  | AMI 135
|  | 9 | J | LHE03 | +V | Helium Level |  | LV/3/H2/J/+V |  | He | New Vent line sensor | AMI 135
|  | 10 | K | LHE03 | -V |  |  | LV/3/H2/K/-V |  |  |  | AMI 135
|  | 11 | L | LHE03 | -I |  |  | LV/3/H2/L/-I |  |  |  | AMI 135
|  | 12 | M | LHE03 | +I |  |  | LV/3/H2/M/+I |  |  |  | AMI 135
|  | 13 | N | [7 empty] | Multiplex to
|  | 14 | P | Common | -I |  |  | LV/1&2/H2/P/-I |  |  | manual switch | 1=+I
|  | 15 | R | [7 empty] | LHE03 used | 6=-V
|  | 16 | S | [7 empty] | when full and | 7=-I
|  | 17 | T | [7 empty] | running | 8=+V
|  | 18 | V | [7 empty] | LHE01 & LHE02
|  | 19 | U | [7 empty] | used during fill
H3 | Ceramic | 1 | A | VTM01 |  | Voltage Tap |  | VT/1/H3/A/MC1+ |  | MC1 + | Current lead 300A |  |  |  |  | 148 | (ohms) |  | 05/18/11 | RMP | 216 | (ohms) |  | 05/18/11 | RMP
| 10 Pin | 2 | B | VTM02 |  | Voltage Tap |  | VT/2/H3/B/MC1- |  | MC1 - | Current lead 300A
| 2.75" CF | 3 | C | VTM03 |  | Voltage Tap |  | VT/3/H3/C/MC2+ |  | MC2 + | Current lead 300A |  |  |  |  | 102 |  |  |  |  | 169
|  | 4 | D | VTM04 |  | Voltage Tap |  | VT/4/H3/D/MC2- |  | MC2 - | Current lead 300A
| (Old J3) | 5 | E | VTM05 |  | Voltage Tap |  | VT/5/H3/E/EC1+ |  | EC1 + | Current lead 300A |  |  |  |  | 118 |  |  |  |  | 184
|  | 6 | F | VTM06 |  | Voltage Tap |  | VT/6/H3/F/EC1-/CC+ |  | EC1 - / CC + | Trimming supply 50A | [5 empty] | 408 |  |  |  |  | 477
|  | 7 | G | VTM07 |  | Voltage Tap |  | VT/7/H3/G/EC2+/CC- |  | EC2 + / CC - | Trimming supply 50A |  |  |  |  | 129 |  |  |  |  | 196
|  | 8 | H | VTM08 |  | Voltage Tap |  | VT/8/H3/H/EC2- |  | EC2 - | Current lead 300A
|  | 9 | J | VTM09 |  | Voltage Tap |  | VT/9/H3/J/CT |  | CT | Centre Tap
|  | 10 | K
V1 | Ceramic | 1 | A | VTL01 |  | Voltage Tap |  | VT/1/V1/A/MC1+/U |  | MC1 + Upper | Current lead 300A 14mm dia
| 19 Pin | 2 | B | VTL02 |  | Voltage Tap |  | VT/2/V1/B/MC1-/U |  | MC1 - Upper | Current lead 300A 14mm dia
| 2.75" CF | 3 | C | VTL03 |  | Voltage Tap |  | VT/3/V1/C/MC2+/U |  | MC2 + Upper | Current lead 300A 14mm dia
|  | 4 | D | VTL04 |  | Voltage Tap |  | VT/4/V1/D/MC2-/U |  | MC2 - Upper | Current lead 300A 14mm dia
| (Old J4) | 5 | E | VTL05 |  | Voltage Tap |  | VT/5/V1/E/EC1+/U |  | EC1 + Upper | Current lead 300A 14mm dia
|  | 6 | F | VTL06 |  | Voltage Tap |  | VT/6/V1/F/EC1-/CC+/U |  | EC1 - / CC + Upper | Trimming Supply lead 50A 11mm dia
|  | 7 | G | VTL07 |  | Voltage Tap |  | VT/7/V1/G/EC2+/CC-/U |  | EC2 + / CC - Upper | Trimming Supply lead 50A 11mm dia
|  | 8 | H | VTL08 |  | Voltage Tap |  | VT/8/V1/H/EC2-/U |  | EC2 - Upper | Current lead 300A 14mm dia
|  | 9 | J | VTL09 |  | Voltage Tap |  | VT/9/V1/J/MC1+/L |  | MC1 + Lower | Current lead 300A 14mm dia
|  | 10 | K | VTL10 |  | Voltage Tap |  | VT/10/V1/K/MC1-/L |  | MC1 - Lower | Current lead 300A 14mm dia
|  | 11 | L | VTL11 |  | Voltage Tap |  | VT/11/V1/L/MC2+/L |  | MC2 + Lower | Current lead 300A 14mm dia
|  | 12 | M | VTL12 |  | Voltage Tap |  | VT/12/V1/M/MC2-/L |  | MC2 - Lower | Current lead 300A 14mm dia
|  | 13 | N | VTL13 |  | Voltage Tap |  | VT/13/V1/N/EC1+/L |  | EC1 + Lower | Current lead 300A 14mm dia
|  | 14 | P | VTL14 |  | Voltage Tap |  | VT/14/V1/P/EC1-/CC+/L |  | EC1 - / CC + Lower | Trimming Supply lead 50A 11mm dia
|  | 15 | R | VTL15 |  | Voltage Tap |  | VT/15/V1/R/EC2+/CC-/L |  | EC2 + / CC - Lower | Trimming Supply lead 50A 11mm dia
|  | 16 | S | VTL16 |  | Voltage Tap |  | VT/16/V1/S/EC2-/L |  | EC2 - Lower | Current lead 300A 14mm dia
|  | 17 | T
|  | 18 | U
|  | 19 | V
V2 | Ceramic | 1 | A | SD01 | +V | Silicon Diode | DT-470-CU-11A | SD/1/V2/A/+V | D6024434 | Single Stage | Copper plate near to HTS leads | Lakeshore 218 | 3 | 5 | 4
| 41 Pin | 2 | B | SD01 | -V |  |  | SD/1/V2/B/-V |  | Cooler plate | Use as an interlock sensor | Lakeshore 218 | 3 | 5 | 16
| 3.375" CF | 3 | C | SD01 | +I |  |  | SD/1/V2/C/+I |  |  |  | Lakeshore 218 | 3 | 5 | 3
|  | 4 | D | SD01 | -I |  |  | SD/1/V2/D/-I |  |  |  | Lakeshore 218 | 3 | 5 | 15
|  | 5 | E | SD16 | +V | Silicon Diode | DT-470-CU-11A | SD/16/V2/E/+V | D6024425 | Cryo cooler 1 |  | Lakeshore 218 | 2 | 2 | 7
| (Old J5) | 6 | F | SD16 | -V |  |  | SD/16/V2/F/-V |  | copper |  | Lakeshore 218 | 2 | 2 | 19
|  | 7 | G | SD16 | +I |  |  | SD/16/V2/G/+I |  | Plate |  | Lakeshore 218 | 2 | 2 | 6
|  | 8 | H | SD16 | -I |  |  | SD/16/V2/H/-I |  |  |  | Lakeshore 218 | 2 | 2 | 18
|  | 9 | J | SD03 | +V | Silicon Diode | DT-470-CU-11A | SD/3/V2/J/+V | D6024432 | Cooler 5 copper |  | Lakeshore 218 | 3 | 2 | 7
|  | 10 | K | SD03 | -V |  |  | SD/3/V2/K/-V |  | plate |  | Lakeshore 218 | 3 | 2 | 19
|  | 11 | L | SD03 | +I |  |  | SD/3/V2/L/+I |  |  |  | Lakeshore 218 | 3 | 2 | 6
|  | 12 | M | SD03 | -I |  |  | SD/3/V2/M/-I |  |  |  | Lakeshore 218 | 3 | 2 | 18
|  | 13 | N | SD27 | +V | Silicon Diode | DT-470-CU-11A | SD/27/V2/N/+V | D6024419 | On the single |  | Lakeshore 218 | 3 | 6 | 7
|  | 14 | P | SD27 | -V |  |  | SD/27/V2/P/-V |  | stage cooler |  | Lakeshore 218 | 3 | 6 | 19
|  | 15 | R | SD27 | +I |  |  | SD/27/V2/R/+I |  | head surface |  | Lakeshore 218 | 3 | 6 | 6
|  | 16 | S | SD27 | -I |  |  | SD/27/V2/S/-I |  |  |  | Lakeshore 218 | 3 | 6 | 18
|  | 17 | T | SD05 | +V | Silicon Diode | DT-470-CU-11A | SD/5/V2/T/+V | D6024424 | Copper Plate |  | Lakeshore 218 | 2 | 3 | 10
|  | 18 | U | SD05 | -V |  |  | SD/5/V2/U/-V |  | cooler 2 |  | Lakeshore 218 | 2 | 3 | 22
|  | 19 | V | SD05 | +I |  |  | SD/5/V2/V/+I |  |  |  | Lakeshore 218 | 2 | 3 | 9
|  | 20 | W | SD05 | -I |  |  | SD/5/V2/W/-I |  |  |  | Lakeshore 218 | 2 | 3 | 21
|  | 21 | X | SD06 | +V | Silicon Diode | DT-470-CU-11A | SD/6/V2/X/+V | D6024437 | Copper Plate |  | Lakeshore 218 | 2 | 4 | 13
|  | 22 | Y | SD06 | -V |  |  | SD/6/V2/Y/-V |  | cooler 3 |  | Lakeshore 218 | 2 | 4 | 25
|  | 23 | Z | SD06 | +I |  |  | SD/6/V2/Z/+I |  |  |  | Lakeshore 218 | 2 | 4 | 12
|  | 24 | a | SD06 | -I |  |  | SD/6/V2/a/-I |  |  |  | Lakeshore 218 | 2 | 4 | 24
|  | 25 | b | SD07 | +V | Silicon Diode | DT-470-CU-11A | SD/7/V2/b/+V | D6024429 | Copper Plate |  | Lakeshore 218 | 3 | 1 | 4
|  | 26 | c | SD07 | -V |  |  | SD/7/V2/c/-V |  | cooler 4 |  | Lakeshore 218 | 3 | 1 | 16
|  | 27 | d | SD07 | +I |  |  | SD/7/V2/d/+I |  |  |  | Lakeshore 218 | 3 | 1 | 3
|  | 28 | e | SD07 | -I |  |  | SD/7/V2/e/-I |  |  |  | Lakeshore 218 | 3 | 1 | 15
|  | 29 | f | SD08 | +V | Silicon Diode | DT-470-CU-11A | SD/8/V2/f/+V | D6024427 | Copper Plate |  | Lakeshore 218 | 3 | 7 | 10
|  | 30 | g | SD08 | -V |  |  | SD/8/V2/g/-V |  | HTS |  | Lakeshore 218 | 3 | 7 | 22
|  | 31 | h | SD08 | +I |  |  | SD/8/V2/h/+I |  | Single stage |  | Lakeshore 218 | 3 | 7 | 9
|  | 32 | i | SD08 | -I |  |  | SD/8/V2/i/-I |  |  |  | Lakeshore 218 | 3 | 7 | 21
|  | 33 | j | CX10 | +V | Cernox | CX-1070-CU | CX/10/V2/j/+V | X66404 | Cold Mass | Change to Cernox from Silicon Diode | Lakeshore 218 | 2 | 6 | 7
|  | 34 | k | CX10 | -V |  |  | CX/10/V2/k/-V |  | Feedthroughs | to make up numbers on monitor 1 | Lakeshore 218 | 2 | 6 | 19
|  | 35 | m | CX10 | +I |  |  | CX/10/V2/m/+I |  |  |  | Lakeshore 218 | 2 | 6 | 6
|  | 36 | n | CX10 | -I |  |  | CX/10/V2/n/-I |  |  |  | Lakeshore 218 | 2 | 6 | 18
|  | 37 | p | SD14 | +V | Silicon Diode | DT-470-CU-11A | SD/14/V2/p/+V | D79371 | Cold Head | Single stage cold head sensor moved from J11 | Lakeshore 218 | 3 | 8 | 13
|  | 38 | q | SD14 | -V |  |  | SD/14/V2/q/-V |  | single stage | on the strap | Lakeshore 218 | 3 | 8 | 25
|  | 39 | r | SD14 | +I |  |  | SD/14/V2/r/+I |  | strap |  | Lakeshore 218 | 3 | 8 | 12
|  | 40 | s | SD14 | -I |  |  | SD/14/V2/s/-I |  |  |  | Lakeshore 218 | 3 | 8 | 24
|  | 41 | t
V3 | Ceramic | 1 | A | SD04 | +V | Silicon Diode | DT-470-CU-11A | SD/4/V3/A/+V | D79362 | Rad Shield |  | Lakeshore 218 | 4 | 1 | 4
| 41 Pin | 2 | B | SD04 | -V |  |  | SD/4/V3/B/-V |  | near MC end |  | Lakeshore 218 | 4 | 1 | 16
| 3.375" CF | 3 | C | SD04 | +I |  |  | SD/4/V3/C/+I |  | upper |  | Lakeshore 218 | 4 | 1 | 3
|  | 4 | D | SD04 | -I |  |  | SD/4/V3/D/-I |  | support |  | Lakeshore 218 | 4 | 1 | 15
|  | 5 | E | SD19 | +V | Silicon Diode |  | SD/19/V3/E/+V | D6026435 | Rad Shield |  | Lakeshore 218 | 4 | 2 | 7
| (Old J6) | 6 | F | SD19 | -V |  | DT-670D-LR | SD/19/V3/F/-V |  |  |  | Lakeshore 218 | 4 | 2 | 19
|  | 7 | G | SD19 | +I |  |  | SD/19/V3/G/+I |  |  |  | Lakeshore 218 | 4 | 2 | 6
|  | 8 | H | SD19 | -I |  |  | SD/19/V3/H/-I |  |  |  | Lakeshore 218 | 4 | 2 | 18
|  | 9 | J | SD20 | +V | Silicon Diode | DT-470-CU-11A | SD/20/V3/J/+V | D80353 | CM Support |  | Lakeshore 218 | 4 | 3 | 10
|  | 10 | K | SD20 | -V |  |  | SD/20/V3/K/-V |  | EC end lower |  | Lakeshore 218 | 4 | 3 | 22
|  | 11 | L | SD20 | +I |  |  | SD/20/V3/L/+I |  | support |  | Lakeshore 218 | 4 | 3 | 9
|  | 12 | M | SD20 | -I |  |  | SD/20/V3/M/-I |  |  |  | Lakeshore 218 | 4 | 3 | 21
|  | 13 | N | SD23 | +V | Silicon Diode |  | SD/23/V3/N/+V | D6026434 | Rad Shield |  | Lakeshore 218 | 4 | 4 | 13
|  | 14 | P | SD23 | -V |  | DT-670D-LR | SD/23/V3/P/-V |  | Inside Bore |  | Lakeshore 218 | 4 | 4 | 25
|  | 15 | R | SD23 | +I |  |  | SD/23/V3/R/+I |  | Matching Coil end |  | Lakeshore 218 | 4 | 4 | 12
|  | 16 | S | SD23 | -I |  |  | SD/23/V3/S/-I |  |  |  | Lakeshore 218 | 4 | 4 | 24
|  | 17 | T | SD24 | +V | Silicon Diode | DT-470-CU-11A | SD/24/V3/T/+V | D79366 | CM Support |  | Lakeshore 218 | 4 | 5 | 4
|  | 18 | U | SD24 | -V |  |  | SD/24/V3/U/-V |  | Matching coil |  | Lakeshore 218 | 4 | 5 | 16
|  | 19 | V | SD24 | +I |  |  | SD/24/V3/V/+I |  | end, upper |  | Lakeshore 218 | 4 | 5 | 3
|  | 20 | W | SD24 | -I |  |  | SD/24/V3/W/-I |  | support |  | Lakeshore 218 | 4 | 5 | 15
|  | 21 | X | SD02 | +V | Silicon Diode | DT-470-CU-11A | SD/2/V3/X/+V | D79370 | Rad Shield |  | Lakeshore 218 | 4 | 6 | 7
|  | 22 | Y | SD02 | -V |  |  | SD/2/V3/Y/-V |  | near vertical |  | Lakeshore 218 | 4 | 6 | 19
|  | 23 | Z | SD02 | +I |  |  | SD/2/V3/Z/+I |  | copper |  | Lakeshore 218 | 4 | 6 | 6
|  | 24 | a | SD02 | -I |  |  | SD/2/V3/a/-I |  | plates |  | Lakeshore 218 | 4 | 6 | 18
|  | 25 | b | SD09 | +V | Silicon Diode | DT-470-CU-11A | SD/9/V3/b/+V | D79374 | Rad Shield |  | Lakeshore 218 | 4 | 7 | 10
|  | 26 | c | SD09 | -V |  |  | SD/9/V3/c/-V |  |  |  | Lakeshore 218 | 4 | 7 | 22
|  | 27 | d | SD09 | +I |  |  | SD/9/V3/d/+I |  |  |  | Lakeshore 218 | 4 | 7 | 9
|  | 28 | e | SD09 | -I |  |  | SD/9/V3/e/-I |  |  |  | Lakeshore 218 | 4 | 7 | 21
|  | 29 | f
|  | 30 | g
|  | 31 | h
|  | 32 | i
|  | 33 | j
|  | 34 | k
|  | 35 | m
|  | 36 | n
|  | 37 | p
|  | 38 | q
|  | 39 | r
|  | 40 | s
|  | 41 | t
V4 | Ceramic | 1 | A | SD25 | +V | Silicon Diode |  | SD/25/V4/A/+V | D6026437 | Centre of |  | Lakeshore 218 | 3 | 3 | 10
| 41 Pin | 2 | B | SD25 | -V |  | DT-670D-LR | SD/25/V4/B/-V |  | rad bore |  | Lakeshore 218 | 3 | 3 | 22
| 3.375" CF | 3 | C | SD25 | +I |  |  | SD/25/V4/C/+I |  | at bottom |  | Lakeshore 218 | 3 | 3 | 9
|  | 4 | D | SD25 | -I |  |  | SD/25/V4/D/-I |  |  |  | Lakeshore 218 | 3 | 3 | 21
|  | 5 | E | SD26 | +V | Silicon Diode | DT-470-CU-11A | SD/26/V4/E/+V | D79369 | Bottom of |  | Lakeshore 218 | 3 | 4 | 13
|  | 6 | F | SD26 | -V |  |  | SD/26/V4/F/-V |  | outer cylinder |  | Lakeshore 218 | 3 | 4 | 25
|  | 7 | G | SD26 | +I |  |  | SD/26/V4/G/+I |  | of rad shield |  | Lakeshore 218 | 3 | 4 | 12
|  | 8 | H | SD26 | -I |  |  | SD/26/V4/H/-I |  |  |  | Lakeshore 218 | 3 | 4 | 24
|  | 9 | J | CX12 | +V | Cernox | CX-1070-CU | CX/12/V4/J/+V | X66443 | Cold Mass | Change to Cernox from Silicon Diode | Lakeshore 218 | 2 | 7 | 10
|  | 10 | K | CX12 | -V |  |  | CX/12/V4/K/-V |  | bottom | to make up numbers on monitor 1 | Lakeshore 218 | 2 | 7 | 22
|  | 11 | L | CX12 | +I |  |  | CX/12/V4/L/+I |  | cover |  | Lakeshore 218 | 2 | 7 | 9
|  | 12 | M | CX12 | -I |  |  | CX/12/V4/M/-I |  |  |  | Lakeshore 218 | 2 | 7 | 21
|  | 13 | N | CX1 | +V | Cernox | CX-1070-CU | CX/1/V4/N/+V | X67217 | Cold Mass |  | Lakeshore 218 | 5 | 5 | 13
|  | 14 | P | CX1 | -V |  |  | CX/1/V4/P/-V |  | Bottom of vessel |  | Lakeshore 218 | 5 | 5 | 25
|  | 15 | R | CX1 | +I |  |  | CX/1/V4/R/+I |  | Matching coil end |  | Lakeshore 218 | 5 | 5 | 12
|  | 16 | S | CX1 | -I |  |  | CX/1/V4/S/-I |  |  |  | Lakeshore 218 | 5 | 5 | 24
|  | 17 | T | CX11 | +V | Cernox | CX-1070-CU | CX/11/V4/T/+V | X66590 | Cold Mass | Change to Cernox from Silicon Diode | Lakeshore 218 | 2 | 8 | 13
|  | 18 | U | CX11 | -V |  |  | CX/11/V4/U/-V |  | Fill tube | to make up numbers on monitor 1 | Lakeshore 218 | 2 | 8 | 25
|  | 19 | V | CX11 | +I |  |  | CX/11/V4/V/+I |  |  |  | Lakeshore 218 | 2 | 8 | 12
|  | 20 | W | CX11 | -I |  |  | CX/11/V4/W/-I |  |  |  | Lakeshore 218 | 2 | 8 | 24
|  | 21 | X | CX2 | +V | Cernox | CX-1070-CU | CX/2/V4/X/+V | X72500 | Cold Mass |  | Lakeshore 218 | 5 | 6 | 4
|  | 22 | Y | CX2 | -V |  |  | CX/2/V4/Y/-V |  | Top of vessel |  | Lakeshore 218 | 5 | 6 | 16
|  | 23 | Z | CX2 | +I |  |  | CX/2/V4/Z/+I |  | End coil end |  | Lakeshore 218 | 5 | 6 | 3
|  | 24 | a | CX2 | -I |  |  | CX/2/V4/a/-I |  |  |  | Lakeshore 218 | 5 | 6 | 15
|  | 25 | b
|  | 26 | c
|  | 27 | d
|  | 28 | e
|  | 29 | f
|  | 30 | g
|  | 31 | h
|  | 32 | i
|  | 33 | n
|  | 34 | p
|  | 35 | r
|  | 36 | s
|  | 37 | n
|  | 38 | p
|  | 39 | r
|  | 40 | s
|  | 41 | t
V5 | Ceramic | 1 | A | HT01 |  | NiFe Heater | 5162 | HT/1/V5/A/+ |  |  | New heater feedthrough
| 10 Pin | 2 | B | HT01 |  | NiFe Heater | 5162 | HT/1/V5/B/- |  |  | Heaters at bottom of He can and Fill tube
| 2.75" CF | 3 | C | HT02 |  | NiFe Heater | 5419 | HT/2/V5/C/+
|  | 4 | D | HT02 |  | NiFe Heater | 5419 | HT/2/V5/D/-
|  | 5 | E
|  | 6 | F
|  | 7 | G
|  | 8 | H
|  | 9 | J
|  | 10 | K
C1 | (Old J7) | 1 | A | SD10 | +V | Silicon Diode | DT-470-CU-11A | SD/10/C1/A/+V | D79372 | 1st Stage | He Cryo Cooler 1 1st Stage | Lakeshore 218 | 1 | 1 | 4
|  | 2 | B | SD10 | -V |  |  | SD/10/C1/B/-V |  |  |  | Lakeshore 218 | 1 | 1 | 16
|  | 3 | C | SD10 | +I |  |  | SD/10/C1/C/+I |  |  |  | Lakeshore 218 | 1 | 1 | 3
|  | 4 | D | SD10 | -I |  |  | SD/10/C1/D/-I |  |  |  | Lakeshore 218 | 1 | 1 | 15
|  | 5 | E | CX05 | +V | Cernox | CX-1070-CU | CX/5/C1/E/+V | X66675 | 2nd Stage | He Cryo Cooler 1 2nd Stage | Lakeshore 218 | 1 | 5 | 4
|  | 6 | F | CX05 | -V |  |  | CX/5/C1/F/-V |  |  |  | Lakeshore 218 | 1 | 5 | 16
|  | 7 | G | CX05 | +I |  |  | CX/5/C1/G/+I |  |  |  | Lakeshore 218 | 1 | 5 | 3
|  | 8 | H | CX05 | -I |  |  | CX/5/C1/H/-I |  |  |  | Lakeshore 218 | 1 | 5 | 15
|  | 9 | J
|  | 10 | K
C2 | (Old J8) | 1 | A | SD11 | +V | Silicon Diode | DT-470-CU-11A | SD/11/C2/A/+V | D80357 | 1st Stage | He Cryo Cooler 2 1st Stage | Lakeshore 218 | 1 | 2 | 7
|  | 2 | B | SD11 | -V |  |  | SD/11/C2/B/-V |  |  |  | Lakeshore 218 | 1 | 2 | 19
|  | 3 | C | SD11 | +I |  |  | SD/11/C2/C/+I |  |  |  | Lakeshore 218 | 1 | 2 | 6
|  | 4 | D | SD11 | -I |  |  | SD/11/C2/D/-I |  |  |  | Lakeshore 218 | 1 | 2 | 18
|  | 5 | E | CX06 | +V | Cernox | CX-1070-CU | CX/6/C2/E/+V | X66681 | 2nd Stage | He Cryo Cooler 2 2nd Stage | Lakeshore 218 | 1 | 6 | 7
|  | 6 | F | CX06 | -V |  |  | CX/6/C2/F/-V |  |  |  | Lakeshore 218 | 1 | 6 | 19
|  | 7 | G | CX06 | +I |  |  | CX/6/C2/G/+I |  |  |  | Lakeshore 218 | 1 | 6 | 6
|  | 8 | H | CX06 | -I |  |  | CX/6/C2/H/-I |  |  |  | Lakeshore 218 | 1 | 6 | 18
|  | 9 | J
|  | 10 | K
C3 | (Old J9) | 1 | A | SD12 | +V | Silicon Diode | DT-470-CU-11A | SD/12/C3/A/+V | D79375 | 1st Stage | He Cryo Cooler 3 1st Stage | Lakeshore 218 | 1 | 3 | 10
|  | 2 | B | SD12 | -V |  |  | SD/12/C3/B/-V |  |  |  | Lakeshore 218 | 1 | 3 | 22
|  | 3 | C | SD12 | +I |  |  | SD/12/C3/C/+I |  |  |  | Lakeshore 218 | 1 | 3 | 9
|  | 4 | D | SD12 | -I |  |  | SD/12/C3/D/-I |  |  |  | Lakeshore 218 | 1 | 3 | 21
|  | 5 | E | CX07 | +V | Cernox | CX-1070-CU | CX/7/C3/E/+V | X66685 | 2nd Stage | He Cryo Cooler 3 2nd Stage | Lakeshore 218 | 1 | 7 | 10
|  | 6 | F | CX07 | -V |  |  | CX/7/C3/F/-V |  |  |  | Lakeshore 218 | 1 | 7 | 22
|  | 7 | G | CX07 | +I |  |  | CX/7/C3/G/+I |  |  |  | Lakeshore 218 | 1 | 7 | 9
|  | 8 | H | CX07 | -I |  |  | CX/7/C3/H/-I |  |  |  | Lakeshore 218 | 1 | 7 | 21
|  | 9 | J
|  | 10 | K
C4 |  | 1 | A | SD13 | +V | Silicon Diode | DT-470-CU-11A | SD/13/C4/A/+V | D80359 | 1st Stage | He Cryo Cooler 4 1st Stage | Lakeshore 218 | 1 | 4 | 13
|  | 2 | B | SD13 | -V |  |  | SD/13/C4/B/-V |  |  |  | Lakeshore 218 | 1 | 4 | 25
|  | 3 | C | SD13 | +I |  |  | SD/13/C4/C/+I |  |  |  | Lakeshore 218 | 1 | 4 | 12
|  | 4 | D | SD13 | -I |  |  | SD/13/C4/D/-I |  |  |  | Lakeshore 218 | 1 | 4 | 24
|  | 5 | E | CX08 | +V | Cernox | CX-1070-CU | CX/8/C4/E/+V | X66687 | 2nd Stage | He Cryo Cooler 4 2nd Stage | Lakeshore 218 | 1 | 8 | 13
|  | 6 | F | CX08 | -V |  |  | CX/8/C4/F/-V |  |  |  | Lakeshore 218 | 1 | 8 | 25
|  | 7 | G | CX08 | +I |  |  | CX/8/C4/G/+I |  |  |  | Lakeshore 218 | 1 | 8 | 12
|  | 8 | H | CX08 | -I |  |  | CX/8/C4/H/-I |  |  |  | Lakeshore 218 | 1 | 8 | 24
|  | 9 | J
|  | 10 | K
C5 |  | 1 | A | SD15 | +V | Silicon Diode | DT-470-CU-11A | SD/15/C5/A/+V | D80358 | 1st Stage | He Cryo Cooler 5 1st Stage | Lakeshore 218 | 2 | 1 | 4
|  | 2 | B | SD15 | -V |  |  | SD/15/C5/B/-V |  |  |  | Lakeshore 218 | 2 | 1 | 16
|  | 3 | C | SD15 | +I |  |  | SD/15/C5/C/+I |  |  |  | Lakeshore 218 | 2 | 1 | 3
|  | 4 | D | SD15 | -I |  |  | SD/15/C5/D/-I |  |  |  | Lakeshore 218 | 2 | 1 | 15
|  | 5 | E | CX09 | +V | Cernox | CX-1070-CU | CX/9/C5/E/+V | X66688 | 2nd Stage | He Cryo Cooler 5 2nd Stage | Lakeshore 218 | 2 | 5 | 4
|  | 6 | F | CX09 | -V |  |  | CX/9/C5/F/-V |  |  |  | Lakeshore 218 | 2 | 5 | 16
|  | 7 | G | CX09 | +I |  |  | CX/9/C5/G/+I |  |  |  | Lakeshore 218 | 2 | 5 | 3
|  | 8 | H | CX09 | -I |  |  | CX/9/C5/H/-I |  |  |  | Lakeshore 218 | 2 | 5 | 15
|  | 9 | J
|  | 10 | K

## Interlock & Alarm Level
Designation | Type | Placement | Comments | Cold Mass above 20K | Radiation shield above 100K | Cooler 1st stage plate above 60k | Match Coil 1 taps at threshold | Match Coil 2 taps at threshold | End Coil 1 taps at threshold | Centre Coil taps at threshold | End Coil 2 taps at threshold | Match 1 HTS taps at threshold | Match 2 HTS taps at threshold | End 1 HTS taps at threshold | Centre HTS taps at threshold | End 2 HTS taps at threshold | Vacuum above Threshold | Vacuum below Threshold | Helium at Alarm Level | Helium at Threshold | Vent Pressure above upper Threshold | Vent Pressure below lower threshold | Power failure (any length of time?) | Cooling water failure ( Threshold to be maximum time) | Compressed air failure (Gate valve normally closed?)
PrTr1 |  | Pressure transducer on vent line | Match to CX12 for feedback
PrTr2 |  | Pressure transducer on fill line | Match to CX11 for feedback
VacG1 |  | Low vacuum gauge
VacG2 |  | High vacuum gauge
GateV |  | Gate valve
Turbo |  | Turbo Pump
Rough |  | Roughing pump
CC1 |  | Cryo Cooler 1
CC2 |  | Cryo Cooler 2
CC3 |  | Cryo Cooler 3
CC4 |  | Cryo Cooler 4
CC5 |  | Cryo Cooler 5
PB1 |  | Power breaker 1
PB2 |  | Power Breaker 2
PB3 |  | Power Breaker 3
C/E 300A |  | Centre / Ends 300A Supply
E1 60A |  | End 1 Trim Supply
E2 60A |  | End 2 trim Supply
M1 300A |  | Match 1 Supply
M2 300A |  | Match 2 Supply
HT01 | NiFe Heater | Fill line heater set
HT02 | NiFe Heater | He can cover heater set
LHE01 | Helium Level | He | Bottom Sensor
LHE02 | Helium Level | He | Top Sensor
LHE03 | Helium Level | He | New Vent line sensor
VTM01 | Voltage Tap | MC1 + | Current lead 300A
VTM02 | Voltage Tap | MC1 - | Current lead 300A
VTM03 | Voltage Tap | MC2 + | Current lead 300A
VTM04 | Voltage Tap | MC2 - | Current lead 300A
VTM05 | Voltage Tap | EC1 + | Current lead 300A
VTM06 | Voltage Tap | EC1 - / CC + | Trimming supply 50A
VTM07 | Voltage Tap | EC2 + / CC - | Trimming supply 50A
VTM08 | Voltage Tap | EC2 - | Current lead 300A
VTM09 | Voltage Tap | CT | Centre Tap
VTL01 | Voltage Tap | MC1 + Upper | Current lead 300A 14mm dia
VTL02 | Voltage Tap | MC1 - Upper | Current lead 300A 14mm dia
VTL03 | Voltage Tap | MC2 + Upper | Current lead 300A 14mm dia
VTL04 | Voltage Tap | MC2 - Upper | Current lead 300A 14mm dia
VTL05 | Voltage Tap | EC1 + Upper | Current lead 300A 14mm dia
VTL06 | Voltage Tap | EC1 - / CC + Upper | Trimming Supply lead 50A 11mm dia
VTL07 | Voltage Tap | EC2 + / CC - Upper | Trimming Supply lead 50A 11mm dia
VTL08 | Voltage Tap | EC2 - Upper | Current lead 300A 14mm dia
VTL09 | Voltage Tap | MC1 + Lower | Current lead 300A 14mm dia
VTL10 | Voltage Tap | MC1 - Lower | Current lead 300A 14mm dia
VTL11 | Voltage Tap | MC2 + Lower | Current lead 300A 14mm dia
VTL12 | Voltage Tap | MC2 - Lower | Current lead 300A 14mm dia
VTL13 | Voltage Tap | EC1 + Lower | Current lead 300A 14mm dia
VTL14 | Voltage Tap | EC1 - / CC + Lower | Trimming Supply lead 50A 11mm dia
VTL15 | Voltage Tap | EC2 + / CC - Lower | Trimming Supply lead 50A 11mm dia
VTL16 | Voltage Tap | EC2 - Lower | Current lead 300A 14mm dia
SD01 | Silicon Diode | Single Stage cooler plate | Copper plate near to HTS leads
SD16 | Silicon Diode | Cryo cooler 1 copper plate
SD03 | Silicon Diode | Cooler 5 copper plate
SD27 | Silicon Diode | On the single stage cooler head surface
SD05 | Silicon Diode | Copper Plate cooler 2
SD06 | Silicon Diode | Copper Plate cooler 3
SD07 | Silicon Diode | Copper Plate cooler 4
SD08 | Silicon Diode | Copper Plate HTS single stage
CX10 | Cernox | Cold Mass feedthroughs
SD14 | Silicon Diode | Cold Head single stage strap
SD04 | Silicon Diode | Rad Shield near MC end upper support
SD19 | Silicon Diode | Rad Shield
SD20 | Silicon Diode | CM Support EC end lower support
SD23 | Silicon Diode | Rad Shield inside bore matching coil end
SD24 | Silicon Diode | CM Support matching coil end upper support
SD02 | Silicon Diode | Rad Shield near vertical copper plates
SD09 | Silicon Diode | Rad Shield
SD25 | Silicon Diode | Centre of of Radiation bore at bottom
SD26 | Silicon Diode | Bottom of outer cylinder of Raditation shield
CX12 | Cernox | Cold Mass bottom cover
SD18 | Silicon Diode | Cold Mass inside bore EC end
CX11 | Cernox | Cold Mass fill tube
SD22 | Silicon Diode | Cold Mass
SD10 | Silicon Diode | 1st Stage | He Cryo Cooler 1 1st Stage
CX05 | Cernox | 2nd Stage | He Cryo Cooler 1 2nd Stage
SD11 | Silicon Diode | 1st Stage | He Cryo Cooler 2 1st Stage
CX06 | Cernox | 2nd Stage | He Cryo Cooler 2 2nd Stage
SD12 | Silicon Diode | 1st Stage | He Cryo Cooler 3 1st Stage
CX07 | Cernox | 2nd Stage | He Cryo Cooler 3 2nd Stage
SD13 | Silicon Diode | 1st Stage | He Cryo Cooler 4 1st Stage
CX08 | Cernox | 2nd Stage | He Cryo Cooler 4 2nd Stage
SD15 | Silicon Diode | 1st Stage | He Cryo Cooler 5 1st Stage
CX09 | Cernox | 2nd Stage | He Cryo Cooler 5 2nd Stage
CX01 | Cernox | He - End Coil 1 | Feedthrough will be on the Fill / Vent tower
CX02 | Cernox | He - Match Coil1

## Monitor Map
Monitor | Input | Type | Feedthrough / Where | Designation | Where | Constant Monitoring | Use for interlock | Interlock comments
Lakeshore 218 1 | 1 | Silicon Diode | J7 Air Side | SD10 | Cryo Cooler 1 1st Stage | Yes
| 2 | Silicon Diode | J8 Air Side | SD11 | Cryo Cooler 2 1st Stage | Yes
| 3 | Silicon Diode | J9 Air Side | SD12 | Cryo Cooler 3 1st Stage | Yes
| 4 | Silicon Diode | J10 Air Side | SD13 | Cryo Cooler 4 1st Stage | Yes
| 5 | Cernox | J7 Air Side | CX05 | Cryo Cooler 1 2nd Stage | Yes
| 6 | Cernox | J8 Air Side | CX06 | Cryo Cooler 2 2nd Stage | Yes
| 7 | Cernox | J9 Air Side | CX07 | Cryo Cooler 3 2nd Stage | Yes
| 8 | Cernox | J10 Air Side | CX08 | Cryo Cooler 4 2nd Stage | Yes
Monitor | Input | Type |  | Designation | Where | Constant Monitoring | Use for interlock | Interlock comments
Lakeshore 218 2 | 1 | Silicon Diode | J11 Air Side | SD10 | Cryo Cooler 5 1st Stage | Yes
| 2 | Silicon Diode | J5 Vac Can | SD16 | Cryo 1 Cooler Plate | Yes
| 3 | Silicon Diode | J5 Vac Can | SD05 | Cryo 2 Cooler plate | Yes
| 4 | Silicon Diode | J5 Vac Can | SD06 | Cryo 3 Cooler Plate | Yes
| 5 | Cernox | J11 Air Side | CX09 | Cryo Cooler 5 2nd Stage | Yes
| 6 | Cernox | J5 Vac Can | CX10 | Cold Mass Feedthroughs | Yes
| 7 | Cernox | J13 Vac Can | CX12 | Cold Mass bottom cover | Yes
| 8 | Cernox | J13 Vac Can | CX11 | Cold Mass He fill line | Yes
Monitor | Input | Type |  | Designation | Where | Constant Monitoring | Use for interlock | Interlock comments
Lakeshore 218 3 | 1 | Silicon Diode | J5 Vac Can | SD07 | Cryo 4 Cooler Plate | Yes
| 2 | Silicon Diode | J5 Vac Can | SD03 | Cryo 5 Cooler Plate | Yes
| 3 | Silicon Diode | J13 Vac Can | SD25 | Centre bottom of Rad shield bore | Yes
| 4 | Silicon Diode | J13 Vac Can | SD26 | Centre bottom of outer rad shield | Yes
| 5 | Silicon Diode | J5 Vac Can | SD01 | Single Stage Cooler plate near HTS leads | Yes
| 6 | Silicon Diode | J5 Vac Can | SD27 | On surface of single stage cooler | Yes
| 7 | Silicon Diode | J5 Vac Can | SD08 | Single Stage Cooler plate near HTS leads | Yes
| 8 | Silicon Diode | J5 Vac Can | SD14 | Single Stage cooler copper strap | Yes
Monitor | Input | Type |  | Designation | Where | Constant Monitoring | Use for interlock | Interlock comments
Lakeshore 218 4 | 1 | Silicon Diode | J6 Vac Can | SD04 | Rad shield near MC coil near support
| 2 | Silicon Diode | J6 Vac Can | SD19 | Inside rad shiled bore EC end
| 3 | Silicon Diode | J6 Vac Can | SD20 | On knuckle of CM support EC end
| 4 | Silicon Diode | J6 Vac Can | SD23 | Inside cold mass bore MC end
| 5 | Silicon Diode | J6 Vac Can | SD24 | On knuckle of CM support MC end
| 6 | Silicon Diode | J6 Vac Can | SD02 | On rad shiled near to vertical copper plates
| 7 | Silicon Diode | J6 Vac Can | SD09 | Rad shiled near to CM support EC end
| 8 | Silicon Diode | J13 Vac Can | SD18 | Inside Cold Mass bore EC end
Monitor | Input | Type |  | Designation | Where | Constant Monitoring | Use for interlock | Interlock comments
Lakeshore 218 5 | 1 | Silicon Diode | J13 Vac Can | SD22 | Inside Cold Mass bore MC end
2
3
4
| 5 | Cernox | He Vessel MC end | CX01 | cold mass MC end
| 6 | Cernox | He Vessel EC end | CX02 | Cold Mass EC end
7
8

## Sensor Labels
Label Colour | Label | Text Colour | Text Font | No Off | Label Colour | Label | Text Colour | Text Font | Number | Label Colour | Label | Text Colour | Text Font | Number | Label Colour | Label | Text Colour | Text Font | Number
White | CX/1/H1/A/+V | Black | Calibri | 5 | White | CX/1/H1/B/-V | Black | Calibri | 5 | White | CX/1/H1/C/+I | Black | Calibri | 5 | White | CX/1/H1/D/-I | Black | Calibri | 5
White | CX/2/H1/E/+V | Black | Calibri | 5 | White | CX/2/H1/F/-V | Black | Calibri | 5 | White | CX/2/H1/G/+I | Black | Calibri | 5 | White | CX/2/H1/H/-I | Black | Calibri | 5
White | LV/1/H2/A/+V | Black | Calibri | 5 | White | LV/1/H2/B/-V | Black | Calibri | 5 | White | LV/1/H2/D/+I | Black | Calibri | 5 | White | LV/3/H2/L/-I | Black | Calibri | 5
White | LV/2/H2/E/+V | Black | Calibri | 5 | White | LV/2/H2/F/-V | Black | Calibri | 5 | White | LV/2/H2/H/+I | Black | Calibri | 5 | White | SD/1/V2/D/-I | Black | Calibri | 5
White | LV/3/H2/J/+V | Black | Calibri | 5 | White | LV/3/H2/K/-V | Black | Calibri | 5 | White | LV/3/H2/M/+I | Black | Calibri | 5 | White | SD/16/V2/H/-I | Black | Calibri | 5
White | SD/1/V2/A/+V | Black | Calibri | 5 | White | SD/1/V2/B/-V | Black | Calibri | 5 | White | SD/1/V2/C/+I | Black | Calibri | 5 | White | SD/3/V2/M/-I | Black | Calibri | 5
White | SD/16/V2/E/+V | Black | Calibri | 5 | White | SD/16/V2/F/-V | Black | Calibri | 5 | White | SD/16/V2/G/+I | Black | Calibri | 5 | White | SD/27/V2/S/-I | Black | Calibri | 5
White | SD/3/V2/J/+V | Black | Calibri | 5 | White | SD/3/V2/K/-V | Black | Calibri | 5 | White | SD/3/V2/L/+I | Black | Calibri | 5 | White | SD/5/V2/W/-I | Black | Calibri | 5
White | SD/27/V2/N/+V | Black | Calibri | 5 | White | SD/27/V2/P/-V | Black | Calibri | 5 | White | SD/27/V2/R/+I | Black | Calibri | 5 | White | SD/6/V2/a/-I | Black | Calibri | 5
White | SD/5/V2/T/+V | Black | Calibri | 5 | White | SD/5/V2/U/-V | Black | Calibri | 5 | White | SD/5/V2/V/+I | Black | Calibri | 5 | White | SD/7/V2/e/-I | Black | Calibri | 5
White | SD/6/V2/X/+V | Black | Calibri | 5 | White | SD/6/V2/Y/-V | Black | Calibri | 5 | White | SD/6/V2/Z/+I | Black | Calibri | 5 | White | SD/8/V2/i/-I | Black | Calibri | 5
White | SD/7/V2/b/+V | Black | Calibri | 5 | White | SD/7/V2/c/-V | Black | Calibri | 5 | White | SD/7/V2/d/+I | Black | Calibri | 5 | White | CX/10/V2/n/-I | Black | Calibri | 5
White | SD/8/V2/f/+V | Black | Calibri | 5 | White | SD/8/V2/g/-V | Black | Calibri | 5 | White | SD/8/V2/h/+I | Black | Calibri | 5 | White | SD/14/V2/s/-I | Black | Calibri | 5
White | CX/10/V2/j/+V | Black | Calibri | 5 | White | CX/10/V2/k/-V | Black | Calibri | 5 | White | CX/10/V2/m/+I | Black | Calibri | 5 | White | SD/4/V3/D/-I | Black | Calibri | 5
White | SD/14/V2/p/+V | Black | Calibri | 5 | White | SD/14/V2/q/-V | Black | Calibri | 5 | White | SD/14/V2/r/+I | Black | Calibri | 5 | White | SD/19/V3/H/-I | Black | Calibri | 5
White | SD/4/V3/A/+V | Black | Calibri | 5 | White | SD/4/V3/B/-V | Black | Calibri | 5 | White | SD/4/V3/C/+I | Black | Calibri | 5 | White | SD/20/V3/M/-I | Black | Calibri | 5
White | SD/19/V3/E/+V | Black | Calibri | 5 | White | SD/19/V3/F/-V | Black | Calibri | 5 | White | SD/19/V3/G/+I | Black | Calibri | 5 | White | SD/23/V3/S/-I | Black | Calibri | 5
White | SD/20/V3/J/+V | Black | Calibri | 5 | White | SD/20/V3/K/-V | Black | Calibri | 5 | White | SD/20/V3/L/+I | Black | Calibri | 5 | White | SD/24/V3/W/-I | Black | Calibri | 5
White | SD/23/V3/N/+V | Black | Calibri | 5 | White | SD/23/V3/P/-V | Black | Calibri | 5 | White | SD/23/V3/R/+I | Black | Calibri | 5 | White | SD/2/V3/a/-I | Black | Calibri | 5
White | SD/24/V3/T/+V | Black | Calibri | 5 | White | SD/24/V3/U/-V | Black | Calibri | 5 | White | SD/24/V3/V/+I | Black | Calibri | 5 | White | SD/9/V3/e/-I | Black | Calibri | 5
White | SD/2/V3/X/+V | Black | Calibri | 5 | White | SD/2/V3/Y/-V | Black | Calibri | 5 | White | SD/2/V3/Z/+I | Black | Calibri | 5 | White | SD/25/V4/D/-I | Black | Calibri | 5
White | SD/9/V3/b/+V | Black | Calibri | 5 | White | SD/9/V3/c/-V | Black | Calibri | 5 | White | SD/9/V3/d/+I | Black | Calibri | 5 | White | SD/26/V4/H/-I | Black | Calibri | 5
White | SD/25/V4/A/+V | Black | Calibri | 5 | White | SD/25/V4/B/-V | Black | Calibri | 5 | White | SD/25/V4/C/+I | Black | Calibri | 5 | White | CX/12/V4/M/-I | Black | Calibri | 5
White | SD/26/V4/E/+V | Black | Calibri | 5 | White | SD/26/V4/F/-V | Black | Calibri | 5 | White | SD/26/V4/G/+I | Black | Calibri | 5 | White | CX/1/V4/S/-I | Black | Calibri | 5
White | CX/12/V4/J/+V | Black | Calibri | 5 | White | CX/12/V4/K/-V | Black | Calibri | 5 | White | CX/12/V4/L/+I | Black | Calibri | 5 | White | CX/11/V4/W/-I | Black | Calibri | 5
White | CX/1/V4/N/+V | Black | Calibri | 5 | White | CX/1/V4/P/-V | Black | Calibri | 5 | White | CX/1/V4/R/+I | Black | Calibri | 5 | White | CX/2/V4/a/-I | Black | Calibri | 5
White | CX/11/V4/T/+V | Black | Calibri | 5 | White | CX/11/V4/U/-V | Black | Calibri | 5 | White | CX/11/V4/V/+I | Black | Calibri | 5 | White | SD/10/C1/D/-I | Black | Calibri | 5
White | CX/2/V4/X/+V | Black | Calibri | 5 | White | CX/2/V4/Y/-V | Black | Calibri | 5 | White | CX/2/V4/Z/+I | Black | Calibri | 5 | White | CX/5/C1/H/-I | Black | Calibri | 5
White | SD/10/C1/A/+V | Black | Calibri | 5 | White | SD/10/C1/B/-V | Black | Calibri | 5 | White | SD/10/C1/C/+I | Black | Calibri | 5 | White | SD/11/C2/D/-I | Black | Calibri | 5
White | CX/5/C1/E/+V | Black | Calibri | 5 | White | CX/5/C1/F/-V | Black | Calibri | 5 | White | CX/5/C1/G/+I | Black | Calibri | 5 | White | CX/6/C2/H/-I | Black | Calibri | 5
White | SD/11/C2/A/+V | Black | Calibri | 5 | White | SD/11/C2/B/-V | Black | Calibri | 5 | White | SD/11/C2/C/+I | Black | Calibri | 5 | White | SD/12/C3/D/-I | Black | Calibri | 5
White | CX/6/C2/E/+V | Black | Calibri | 5 | White | CX/6/C2/F/-V | Black | Calibri | 5 | White | CX/6/C2/G/+I | Black | Calibri | 5 | White | CX/7/C3/H/-I | Black | Calibri | 5
White | SD/12/C3/A/+V | Black | Calibri | 5 | White | SD/12/C3/B/-V | Black | Calibri | 5 | White | SD/12/C3/C/+I | Black | Calibri | 5 | White | SD/13/C4/D/-I | Black | Calibri | 5
White | CX/7/C3/E/+V | Black | Calibri | 5 | White | CX/7/C3/F/-V | Black | Calibri | 5 | White | CX/7/C3/G/+I | Black | Calibri | 5 | White | CX/8/C4/H/-I | Black | Calibri | 5
White | SD/13/C4/A/+V | Black | Calibri | 5 | White | SD/13/C4/B/-V | Black | Calibri | 5 | White | SD/13/C4/C/+I | Black | Calibri | 5 | White | SD/15/C5/D/-I | Black | Calibri | 5
White | CX/8/C4/E/+V | Black | Calibri | 5 | White | CX/8/C4/F/-V | Black | Calibri | 5 | White | CX/8/C4/G/+I | Black | Calibri | 5 | White | CX/9/C5/H/-I | Black | Calibri | 5
White | SD/15/C5/A/+V | Black | Calibri | 5 | White | SD/15/C5/B/-V | Black | Calibri | 5 | White | SD/15/C5/C/+I | Black | Calibri | 5 | White | LV/1&2/H2/P/-I | Black | Calibri | 5
White | CX/9/C5/E/+V | Black | Calibri | 5 | White | CX/9/C5/F/-V | Black | Calibri | 5 | White | CX/9/C5/G/+I | Black | Calibri | 5
Label Colour | Label | Text Colour | Text Font | Number | Label Colour | Label | Text Colour | Text Font | Number | Label Colour | Label | Text Colour | Text Font | Number
White | VT/1/H3/A/MC1+ | Black | Calibri | 5 | White | HT/1/V5/A/+ | Black | Calibri | 5 | White | HT/1/V5/B/- | Black | Calibri | 5
White | VT/2/H3/B/MC1- | Black | Calibri | 5 | White | HT/2/V5/C/+ | Black | Calibri | 5 | White | HT/2/V5/D/- | Black | Calibri | 5
White | VT/3/H3/C/MC2+ | Black | Calibri | 5
White | VT/4/H3/D/MC2- | Black | Calibri | 5
White | VT/5/H3/E/EC1+ | Black | Calibri | 5
White | VT/6/H3/F/EC1-/CC+ | Black | Calibri | 5 |  | Sensor Map Key
White | VT/7/H3/G/EC2+/CC- | Black | Calibri | 5
White | VT/8/H3/H/EC2- | Black | Calibri | 5 |  | Type |  |  | Voltage Tap labels
White | VT/9/H3/J/CT | Black | Calibri | 5 |  | Silicon Diode Sensor | SD |  | Matching Coil | MC
White | VT/1/V1/A/MC1+/U | Black | Calibri | 5 |  | Cernox Sensor | CX |  | End Coil | EC
White | VT/2/V1/B/MC1-/U | Black | Calibri | 5 |  | Voltage Tap | VT |  | Centre Coil | CC
White | VT/3/V1/C/MC2+/U | Black | Calibri | 5 |  | Heater | HT |  | Upper Tap | U
White | VT/4/V1/D/MC2-/U | Black | Calibri | 5 |  | Level Sensor | LV |  | Lower Tap | L
White | VT/5/V1/E/EC1+/U | Black | Calibri | 5
White | VT/6/V1/F/EC1-/CC+/U | Black | Calibri | 5 |  | Feedthrough Designation | [5 empty] | VT/7/V1/G/EC2+/CC-/U |  |  |  |  | SD/15/C5/B/-V
White | VT/7/V1/G/EC2+/CC-/U | Black | Calibri | 5 |  | Helium Can | H
White | VT/8/V1/H/EC2-/U | Black | Calibri | 5 |  | Vacuum Can / Tower | V
White | VT/9/V1/J/MC1+/L | Black | Calibri | 5 |  | Cooler (External) | C
White | VT/10/V1/K/MC1-/L | Black | Calibri | 5
White | VT/11/V1/L/MC2+/L | Black | Calibri | 5
White | VT/12/V1/M/MC2-/L | Black | Calibri | 5
White | VT/13/V1/N/EC1+/L | Black | Calibri | 5
White | VT/14/V1/P/EC1-/CC+/L | Black | Calibri | 5
White | VT/15/V1/R/EC2+/CC-/L | Black | Calibri | 5
White | VT/16/V1/S/EC2-/L | Black | Calibri | 5
Type
Type Number
Feedthrough Pin
On Feedthrough
Coil
Upper / Lower
Type
Feedthrough Pin
Type Number
On Feedthrough
Polarity

## Cable Map
Re-use old cables
Old J4 | V2 |  | Temp Sensors |  |  | Old J3 | V3 |  | Temp Sensors |  |  | Old J1 | V1 |  |  | Old J2 | C1 | Temp Sensors |  |  | Old J8 | C4 | Temp Sensors |  |  | NEW | V4 |  | Temp Sensors |  |  | NEW | V5
Pin | Cable Colour | Bundle | Monitor | Input | Pin | Pin | Cable Colour | Bundle | Monitor | Input | Pin | Pin | Cable Colour |  | Voltage Tap | Pin | Cable Colour | Monitor | Input | Pin | Pin | Cable Colour | Monitor | Input | Pin | Pin | Cable Colour | Bundle | Monitor | Input | Pin | Pin | Cable Colour | Heater
A | Brown | Black cable bundle | 3 | 5 | 4 | A | Brown | Black cable bundle | 4 | 1 | 4 | A | Brown | Pair | VTL01 | A | Black | 1 | 1 | 4 | A | Black | 1 | 4 | 13 | A | Brown | Black cable bundle | 3 | 3 | 10 | A | Black | HT01
B | Black |  | 3 | 5 | 16 | B | Black |  | 4 | 1 | 16 | B | Black |  | VTL02 | B | Brown | 1 | 1 | 16 | B | Brown | 1 | 4 | 25 | B | Black |  | 3 | 3 | 22 | B | Brown | HT01
C | Red |  | 3 | 5 | 3 | C | Red |  | 4 | 1 | 3 | C | Red | Pair | VTL03 | C | Red | 1 | 1 | 3 | C | Red | 1 | 4 | 12 | C | Red |  | 3 | 3 | 9 | C | Red | HT02
D | Orange |  | 3 | 5 | 15 | D | Orange |  | 4 | 1 | 15 | D | Black |  | VTL04 | D | Orange | 1 | 1 | 15 | D | Orange | 1 | 4 | 24 | D | Orange |  | 3 | 3 | 21 | D | Orange | HT02
E | Yellow |  | 2 | 2 | 7 | E | Yellow |  | 4 | 2 | 7 | E | Orange | Pair | VTL05 | E | Yellow | 1 | 5 | 4 | E | Yellow | 1 | 8 | 13 | E | Yellow |  | 3 | 4 | 13 | E | Yellow
F | Green |  | 2 | 2 | 19 | F | Green |  | 4 | 2 | 19 | F | Black |  | VTL06 | F | Green | 1 | 5 | 16 | F | Green | 1 | 8 | 25 | F | Green |  | 3 | 4 | 25 | F | Green
G | Blue |  | 2 | 2 | 6 | G | Blue |  | 4 | 2 | 6 | G | Yellow | Pair | VTL07 | G | Blue | 1 | 5 | 3 | G | Blue | 1 | 8 | 12 | G | Blue |  | 3 | 4 | 12 | G | Blue
H | Purple |  | 2 | 2 | 18 | H | Purple |  | 4 | 2 | 18 | H | Black |  | VTL08 | H | Purple | 1 | 5 | 15 | H | Purple | 1 | 8 | 24 | H | Purple |  | 3 | 4 | 24 | H | Purple
J | Grey |  | 3 | 2 | 7 | J | Grey |  | 4 | 3 | 10 | J | Green | Pair | VTL09 | J | Grey |  |  |  | J | Grey |  |  |  | J | Grey |  | 2 | 7 | 10 | J | Grey
K | White |  | 3 | 2 | 19 | K | White |  | 4 | 3 | 22 | K | Black |  | VTL10 | K | White |  |  |  | K | White |  |  |  | K | White |  | 2 | 7 | 22 | K | White
L | Brown | Pair | 3 | 2 | 6 | L | Brown | Pair | 4 | 3 | 9 | L | Blue | Pair | VTL11 | [10 empty] | L | Brown | Pair | 2 | 7 | 9
M | Black |  | 3 | 2 | 18 | M | Black |  | 4 | 3 | 21 | M | Black |  | VTL12 | Old J7 | C2 |  |  |  | NEW | C5 |  |  |  | M | Black |  | 2 | 7 | 21 | NEW | H2
N | Red | Pair | 3 | 6 | 7 | N | Red | Pair | 4 | 4 | 13 | N | Red | Pair | VTL13 | Pin | Cable Colour |  |  |  | Pin | Cable Colour |  |  |  | N | Red | Pair | 5 | 5 | 13 | Pin | Cable Colour |  | Level sensor | Polarity
P | Black |  | 3 | 6 | 19 | P | Black |  | 4 | 4 | 25 | P | Brown |  | VTL14 | A | Black | 1 | 2 | 7 | A | Black | 2 | 1 | 4 | P | Black |  | 5 | 5 | 25 | A | Brown | Pair | LHE01 | +V
R | Orange | Pair | 3 | 6 | 6 | R | Orange | Pair | 4 | 4 | 12 | R | Red | Yellow Pair | VTL15 | B | Brown | 1 | 2 | 19 | B | Brown | 2 | 1 | 16 | R | Orange | Pair | 5 | 5 | 12 | B | Black |  | LHE01 | -V
S | Black |  | 3 | 6 | 18 | S | Black |  | 4 | 4 | 24 | S | Red | Pair | VTL16 | C | Red | 1 | 2 | 6 | C | Red | 2 | 1 | 3 | S | Black |  | 5 | 5 | 24 | C | Red | Pair
T | Yellow | Pair | 2 | 3 | 10 | T | Yellow | Pair | 4 | 5 | 4 | T | Green |  |  | D | Orange | 1 | 2 | 18 | D | Orange | 2 | 1 | 15 | T | Yellow | Pair | 2 | 8 | 13 | D | Black |  | LHE01 | +I
U | Black |  | 2 | 3 | 22 | U | Black |  | 4 | 5 | 16 | U | Red | Pair |  | E | Yellow | 1 | 6 | 7 | E | Yellow | 2 | 5 | 4 | U | Black |  | 2 | 8 | 25 | E | Orange | Pair | LHE02 | +V
V | Green | Pair | 2 | 3 | 9 | V | Green | Pair | 4 | 5 | 3 | V | Blue |  |  | F | Green | 1 | 6 | 19 | F | Green | 2 | 5 | 16 | V | Green | Pair | 2 | 8 | 12 | F | Black |  | LHE02 | -V
W | Black |  | 2 | 3 | 21 | W | Black |  | 4 | 5 | 15 |  |  |  |  | G | Blue | 1 | 6 | 6 | G | Blue | 2 | 5 | 3 | W | Black |  | 2 | 8 | 24 | G | Yellow | Pair
X | Blue | Pair | 2 | 4 | 13 | X | Blue | Pair | 4 | 6 | 7 | No Old Label |  | H3 |  | H | Purple | 1 | 6 | 18 | H | Purple | 2 | 5 | 15 | X | Blue | Pair | 5 | 6 | 4 | H | Black |  | LHE02 | +I
Y | Black |  | 2 | 4 | 25 | Y | Black |  | 4 | 6 | 19 | Pin | Cable Colour |  |  | J | Grey |  |  |  | J | Grey |  |  |  | Y | Black |  | 5 | 6 | 16 | J | Green | Pair | LHE03 | +V
Z | White | Pair | 2 | 4 | 12 | Z | White | Pair | 4 | 6 | 6 | A | Brown | Pair | VTM01 | K | White |  |  |  | K | White |  |  |  | Z | White | Pair | 5 | 6 | 3 | K | Black |  | LHE03 | -V
a | Black |  | 2 | 4 | 24 | a | Black |  | 4 | 6 | 18 | B | Black |  | VTM02 | [10 empty] | a | Black |  | 5 | 6 | 15 | L | Blue | Pair | LHE03 | -I
b | Red | Pair | 3 | 1 | 4 | b | Red | Pair | 4 | 7 | 10 | C | Red | Pair | VTM03 | Old J6 | C3 |  |  |  | NEW | H1 |  |  |  | b | Red | Pair |  |  |  | M | Black |  | LHE03 | +I
c | Brown |  | 3 | 1 | 16 | c | Brown |  | 4 | 7 | 22 | D | Black |  | VTM04 | Pin | Cable Colour |  |  |  | Pin | Cable Colour |  |  |  | c | Brown |  |  |  |  | N | Red | Pair
d | Red | Pair | 3 | 1 | 3 | d | Red | Pair | 4 | 7 | 9 | E | Orange | Pair | VTM05 | A | Black | 1 | 3 | 10 | A | Black | 5 | 7 | 4 | d | Red | Pair |  |  |  | P | Brown |  | Common | -I
e | Orange |  | 3 | 1 | 15 | e | Orange |  | 4 | 7 | 21 | F | Black |  | VTM06 | B | Brown | 1 | 3 | 22 | B | Brown | 5 | 7 | 16 | e | Orange |  |  |  |  | R | Red | Yellow Pair
f | Red | Pair | 3 | 7 | 10 | f | Red | Pair |  |  |  | G | Yellow | Pair | VTM07 | C | Red | 1 | 3 | 9 | C | Red | 5 | 7 | 3 | f | Red | Pair |  |  |  | S | Red | Pair
g | Yellow |  | 3 | 7 | 22 | g | Yellow |  |  |  |  | H | Black |  | VTM08 | D | Orange | 1 | 3 | 21 | D | Orange | 5 | 7 | 15 | g | Yellow |  |  |  |  | T | Green
h | Red | Pair | 3 | 7 | 9 | h | Red | Pair |  |  |  | J | Green | Pair | VTM09 | E | Yellow | 1 | 7 | 10 | E | Yellow | 5 | 8 | 7 | h | Red | Pair |  |  |  | U | Red | Pair
i | Green |  | 3 | 7 | 21 | i | Green |  |  |  |  | K | Black |  |  | F | Green | 1 | 7 | 22 | F | Green | 5 | 8 | 19 | i | Green |  |  |  |  | V | Blue
j | Red | Pair | 2 | 6 | 7 | j | Red | Pair |  |  |  | L | Blue | Pair |  | G | Blue | 1 | 7 | 9 | G | Blue | 5 | 8 | 6 | j | Red | Pair
K | Blue |  | 2 | 6 | 19 | K | Blue |  |  |  |  | M | Black |  |  | H | Purple | 1 | 7 | 21 | H | Purple | 5 | 8 | 18 | K | Blue
m | Red | Pair | 2 | 6 | 6 | m | Red | Pair |  |  |  | N |  |  |  | J | Grey |  |  |  | J | Grey |  |  |  | m | Red | Pair
n | White |  | 2 | 6 | 18 | n | White |  |  |  |  | P |  |  |  | K | White |  |  |  | K | White |  |  |  | n | White
p | Blue | Pair | 3 | 8 | 13 | p | Blue | Pair |  |  |  | R | [13 empty] | p | Blue | Pair
q | Green |  | 3 | 8 | 25 | q | Green |  |  |  |  | S | [13 empty] | q | Green
r | White | Pair | 3 | 8 | 12 | r | White | Pair |  |  |  | T | [13 empty] | r | White | Pair
s | Green |  | 3 | 8 | 24 | s | Green |  |  |  |  | U | [13 empty] | s | Green
t | No cable |  |  |  |  | t | No cable |  |  |  |  | V | [13 empty] | t | No cable

## Level Sensor
Switch pin
H2 | 1 | A | LHE01 | +V | 10
| 2 | B | LHE01 | -V | 4
| 3 | C
| 4 | D | LHE01 | +I | 1
| 5 | E | LHE02 | +V | 11
| 6 | F | LHE02 | -V | 5
| 7 | G
| 8 | H | LHE02 | +I | 2 | I+
| 9 | J | LHE03 | +V | 12 | V-
| 10 | K | LHE03 | -V | 6 | I- | 1 |  |  | Single AMI 135 Monitor for 3 sensors
| 11 | L | LHE03 | -I | 9 | V+ | 4 | A
| 12 | M | LHE03 | +I | 3 |  | 7 |  | 1
| 13 | N | 0 | 0 |  |  | 10 |  | 2
| 14 | P | Common | -I | 7 & 8 | I+ | 2 | B | 3
| 15 | R | 0 | 0 |  | V- | 5 |  | 4
| 16 | S | 0 | 0 |  | I- | 8 |  | 5
| 17 | T | 0 | 0 |  | V+ | 11 | C | 6
| 18 | V | 0 | 0 |  |  | 3 |  | 7
| 19 | U | 0 | 0 |  |  | 6 |  | 8
[6 empty] | I+ | 9 | D | 9
[6 empty] | V- | 12
AMI 135 Sensor Map | [5 empty] | I-
Pin | Function |  |  |  |  | V+
1 | Sensor I+ | (Red)
2 | Not Used
3 | Not Used
4 | Not Used
5 | Not Used
6 | Sensor V- | (Yellow)
7 | Sensor I- | (Black)
8 | Sensor V+ | (Blue)
9 | Not Used
Manual Rotary Switch
AMI 135
Lv1
Lv2
Lv3

## Operational Modes
Operational Modes, monitoring and interlocks for inclusion into the control system.
Evacuation, Outgassing of the MLI and (Manual operation of the vacuum pumping system)
Interlock |  | Check by | Monitoring |  | Check by
1 | Level gauages confirmed to be OFF | System | 1 | Pirani gauge - Vacuum chamber | System
2 |  |  | 2 | Pirani Gauge - Helium circuit | System
3 |  |  | 3 | Ion Gauge - Vacuum chamber | System
4 |  |  | 4
5 |  |  | 5
6 |  |  | 6
7 |  |  | 7
8 |  |  | 8
9 |  |  | 9
10 |  |  | 10
Purging of the Helium system (Manual operation for the purging gas)
Interlock |  | Check by | Monitoring |  | Check by
1 | Level gauages confirmed to be OFF | System | 1 | Pressure Gauge - Vacuum chamber | System
2 | Pumping system to be powered OFF | Operator | 2 | Pressure Gauge - Helium Circuit | System
3 | Gate Valve CLOSED | Operator | 3
4 | Pirani gauges OFF | Operator | 4
5 | Ion Gauges OFF | Operator | 5
6 |  |  | 6
7 |  |  | 7
8 |  |  | 8
9 |  |  | 9
10 |  |  | 10
Cooldown by LN2 ( Manual operation using FermiLab system)
Interlock |  | Check by | Monitoring |  | Check by
1 | Level gauages confirmed to be OFF | System | 1 | Cold mass temp sensors | System
2 | Sufficient vacuum achieved in He circuit | Operator | 2 | Radiation shield temp sensors | System
3 |  |  | 3 | Cryo cooler temp sensors | System
4 |  |  | 4 | 1st stage temp sensors | System
5 |  |  | 5 | Pressure Gauge on Helium circuit | System
6 |  |  | 6 | Flow meter | Operator
7 |  |  | 7 | Ion Gauge on Vacuum chamber | Operator
8 |  |  | 8
9 |  |  | 9
10 |  |  | 10
Evacuation of Nitrogen from the Helium Circuit (Manual operation with control of heaters through system)
Interlock |  | Check by | Monitoring |  | Check by | Control system operation |  | Input by
1 | Level gauages confirmed to be OFF | System | 1 | Cold mass temp sensors | System | 1 | Power input to He line heaters | Operator
2 | Sensors CX1 and CX2 at 80k | Operator | 2 | Radiation shield temp sensors | System | 2 | Power input to He vessel heaters | Operator
3 |  |  | 3 | Cryo cooler temp sensors | System | 3
4 |  |  | 4 | 1st stage temp sensors | System | 4
5 |  |  | 5 | Pressure Gauge on Helium circuit | System | 5
6 |  |  | 6 | Ion Gauge on Vacuum chamber | Operator | 6
7 |  |  | 7 | Pirani Gauge on the Helium Circuit | Operator | 7
8 |  |  | 8 |  |  | 8
9 |  |  | 9 |  |  | 9
10 |  |  | 10 |  |  | 10
Cooldown by LHe ( Manual operation using FermiLab system)
Interlock |  | Check by | Monitoring |  | Check by | Control system operation |  | Input by
1 | Level gauages confirmed to be OFF | System | 1 | Cold mass temp sensors | System | 1
2 | Sensors CX1 and CX2 at 80k | Operator | 2 | Radiation shield temp sensors | System | 2
3 |  |  | 3 | Cryo cooler temp sensors | System | 3
4 |  |  | 4 | 1st stage temp sensors | System | 4
5 |  |  | 5 | Pressure Gauge on Helium circuit | System | 5
6 |  |  | 6 | Ion Gauge on Vacuum chamber | Operator | 6
7 |  |  | 7 | Pirani Gauge on the Helium Circuit | Operator | 7
8 |  |  | 8 |  |  | 8
9 |  |  | 9 |  |  | 9
10 |  |  | 10 |  |  | 10